## Electrical Parameters
Parameter | Description | Byte Size | R/W | Type | MODBUS Code | MODBUS Code With Offset | Supported Device
PRM_CODE_PQ_COMPLIANCE_NAME | Compliance name (EN 50160 or any other compliance standard) | 40 | R | String | 543 | 30544 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_DESCRIPTION | Compliance description | 80 | R | String | 563 | 30564 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_STATUS | OK / Fail / NA status of the compliance currently running | 4 | R | 2 Bytes Integer | 603 | 30604 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_START_TIME | Time since Compliance activation (time steps are at 100 nsec) | 8 | R | 8 Bytes FileTime | 605 | 30606 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_LAST_TIME | The last time (at 100 nsec steps), if running this is the current time stamp if stopped than it is the last time processed | 8 | R | 8 Bytes FileTime | 609 | 30610 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_AGGREGATE_ON_TIME | Amount of time (delta file time format) the compliance window test is so far (excluding the time the device was off) | 8 | R | 8 Bytes FileTime | 613 | 30614 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_AGGREGATE_OFF_TIME | Amount of time (delta file time format) the device was down inside the window | 8 | R | 8 Bytes FileTime | 617 | 30618 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_PRCNTG_TIME_OK | Percentage from the entire time (including off time) in which this compliance parameters was ok. | 4 | R | 4 Bytes Float | 621 | 30622 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_PRCNTG_TIME_FAIL | Percentage from the entire time (including off time) in which this the compliance parameter was not ok. | 4 | R | 4 Bytes Float | 623 | 30624 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_PRCNTG_TIME_NA | Percentage from the entire time (including off time) in which this compliance parameter was not evaluated (due to device off time) | 4 | R | 4 Bytes Float | 625 | 30626 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_NUM_PARAMS | Number of compliance parameters under check | 4 | R | 4 Bytes Integer | 627 | 30628 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM1_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 383 | 30384 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM1_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 407 | 30408 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM2_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 385 | 30386 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM2_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 408 | 30409 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM3_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 387 | 30388 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM3_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 409 | 30410 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM4_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 389 | 30390 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM4_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 410 | 30411 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM5_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 391 | 30392 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM5_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 411 | 30412 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM6_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 393 | 30394 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM6_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 412 | 30413 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM7_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 395 | 30396 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM7_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 413 | 30414 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM8_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 397 | 30398 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM8_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 414 | 30415 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM9_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 399 | 30400 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM9_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 415 | 30416 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM10_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 401 | 30402 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM10_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 416 | 30417 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM11_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 403 | 30404 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM11_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 417 | 30418 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM12_PRCNTG_TIME_OK |  | 4 | R | 4 Bytes Float | 405 | 30406 | Power Quality Analyzer G4400
PRM_CODE_PQ_PRM12_TOT_NUM_EVENTS |  | 2 | R | 4 Bytes Integer | 418 | 30419 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V1 |  | 2 | R | 2 Bytes Integer | 2716 | 32717 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V2 |  | 2 | R | 2 Bytes Integer | 2717 | 32718 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V3 |  | 2 | R | 2 Bytes Integer | 2718 | 32719 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V12 |  | 2 | R | 2 Bytes Integer | 2719 | 32720 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V23 |  | 2 | R | 2 Bytes Integer | 2720 | 32721 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_DIP_V31 |  | 2 | R | 2 Bytes Integer | 2721 | 32722 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V1 |  | 2 | R | 2 Bytes Integer | 2722 | 32723 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V2 |  | 2 | R | 2 Bytes Integer | 2723 | 32724 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V3 |  | 2 | R | 2 Bytes Integer | 2724 | 32725 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V12 |  | 2 | R | 2 Bytes Integer | 2725 | 32726 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V23 |  | 2 | R | 2 Bytes Integer | 2726 | 32727 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_OVERVOLTAGE_V31 |  | 2 | R | 2 Bytes Integer | 2727 | 32728 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V1 |  | 2 | R | 2 Bytes Integer | 2728 | 32729 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V2 |  | 2 | R | 2 Bytes Integer | 2729 | 32730 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V3 |  | 2 | R | 2 Bytes Integer | 2730 | 32731 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V12 |  | 2 | R | 2 Bytes Integer | 2731 | 32732 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V23 |  | 2 | R | 2 Bytes Integer | 2732 | 32733 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_SHORTINT_V31 |  | 2 | R | 2 Bytes Integer | 2733 | 32734 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V1 |  | 2 | R | 2 Bytes Integer | 2734 | 32735 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V2 |  | 2 | R | 2 Bytes Integer | 2735 | 32736 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V3 |  | 2 | R | 2 Bytes Integer | 2736 | 32737 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V12 |  | 2 | R | 2 Bytes Integer | 2737 | 32738 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V23 |  | 2 | R | 2 Bytes Integer | 2738 | 32739 | Power Quality Analyzer G4400
PRM_CODE_PQ_COUNTER_LONGINT_V31 |  | 2 | R | 2 Bytes Integer | 2739 | 32740 | Power Quality Analyzer G4400
PRM_CODE_PSU_TEMPERATURE | power supply temperature (read every second) | 4 | R | 4 Bytes Float | 537 | 30538 | Power Quality Analyzer G4400
PRM_CODE_DSP_FREQUENCY_FAST | The measured frequency , calculated every cycle | 4 | R | 4 Bytes Float | 999 | 31000 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_1 | Line 1 active power (W), calculate every cycle | 4 | R | 4 Bytes Float | 1025 | 31026 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_2 | Line 2 active power (W), calculate every cycle | 4 | R | 4 Bytes Float | 1027 | 31028 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_3 | Line 3 active power (W), calculate every cycle | 4 | R | 4 Bytes Float | 1029 | 31030 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_SUM | Total Active Power (W), calculate every cycle | 4 | R | 4 Bytes Float | 1037 | 31038 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_1 | Line 1 Reactive power (VAr) calculate every cycle | 4 | R | 4 Bytes Float | 1041 | 31042 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_2 | Line 2 Reactive power (VAr) calculate every cycle | 4 | R | 4 Bytes Float | 1043 | 31044 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_3 | Line 3 Reactive power (VAr) calculate every cycle | 4 | R | 4 Bytes Float | 1045 | 31046 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_SUM | Reactive power - Total (VAr) calculate every cycle | 4 | R | 4 Bytes Float | 1053 | 31054 | Power Quality Analyzer G4400
PRM_CODE_APPARENT_POWER_MAIN_1 | Line 1 apparent power (VA) calculate every cycle | 4 | R | 4 Bytes Float | 1057 | 31058 | Power Quality Analyzer G4400
PRM_CODE_APPARENT_POWER_MAIN_2 | Line 2 apparent power (VA) calculate every cycle | 4 | R | 4 Bytes Float | 1059 | 31060 | Power Quality Analyzer G4400
PRM_CODE_APPARENT_POWER_MAIN_3 | Line 3 apparent power (VA) calculate every cycle | 4 | R | 4 Bytes Float | 1061 | 31062 | Power Quality Analyzer G4400
PRM_CODE_APPARENT_POWER_SUM | Apparent Power Total (VA) calculate every cycle | 4 | R | 4 Bytes Float | 1069 | 31070 | Power Quality Analyzer G4400
PRM_CODE_V1_RMS | RMS Voltage V1, calculate every cycle | 4 | R | 4 Bytes Float | 1103 | 31104 | Power Quality Analyzer G4400
PRM_CODE_V2_RMS | RMS Voltage V2, calculate every cycle | 4 | R | 4 Bytes Float | 1105 | 31106 | Power Quality Analyzer G4400
PRM_CODE_V3_RMS | RMS Voltage V3, calculate every cycle | 4 | R | 4 Bytes Float | 1107 | 31108 | Power Quality Analyzer G4400
PRM_CODE_V12_RMS | RMS Voltage V12, calculate every cycle | 4 | R | 4 Bytes Float | 1109 | 31110 | Power Quality Analyzer G4400
PRM_CODE_V23_RMS | RMS Voltage V23, calculate every cycle | 4 | R | 4 Bytes Float | 1111 | 31112 | Power Quality Analyzer G4400
PRM_CODE_V31_RMS | RMS Voltage V31, calculate every cycle | 4 | R | 4 Bytes Float | 1113 | 31114 | Power Quality Analyzer G4400
PRM_CODE_I1_RMS | RMS Voltage I1, calculate every cycle | 4 | R | 4 Bytes Float | 1117 | 31118 | Power Quality Analyzer G4400
PRM_CODE_I2_RMS | RMS Voltage I2, calculate every cycle | 4 | R | 4 Bytes Float | 1119 | 31120 | Power Quality Analyzer G4400
PRM_CODE_I3_RMS | RMS current I3, calculate every cycle | 4 | R | 4 Bytes Float | 1121 | 31122 | Power Quality Analyzer G4400
PRM_CODE_V1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Voltage | 4 | R | 4 Bytes Float | 1123 | 31124 | Power Quality Analyzer G4400
PRM_CODE_V12_23_31_RMS_AVERAGE | Cycle by cycle Average RMS phase to phase Voltage | 4 | R | 4 Bytes Float | 1125 | 31126 | Power Quality Analyzer G4400
PRM_CODE_I1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Current | 4 | R | 4 Bytes Float | 1127 | 31128 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_INTERNAL | Internal DSP temperature, read every cycle | 4 | R | 4 Bytes Float | 1159 | 31160 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_EXTERNAL | External DSP temperature (PT100 READING) , read every cycle | 4 | R | 4 Bytes Float | 1161 | 31162 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_2SEC | PST (short term flicker) - L1 - calculated over 2 seconds | 4 | R | 4 Bytes Float | 1659 | 31660 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_2SEC | PST (short term flicker) - L1 - calculated over 2 seconds | 4 | R | 4 Bytes Float | 1661 | 31662 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_2SEC | PST (short term flicker) - L1 - calculated over 2 seconds | 4 | R | 4 Bytes Float | 1663 | 31664 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_2SEC | PST (short term flicker)- L12 - calculated over 2 seconds | 4 | R | 4 Bytes Float | 1665 | 31666 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_2SEC | PST (short term flicker)- L23 -calculated over 2 seconds | 4 | R | 4 Bytes Float | 1667 | 31668 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_2SEC | PST (short term flicker)- L31 - calculated over 2 seconds | 4 | R | 4 Bytes Float | 1669 | 31670 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10SEC | PST (short term flicker) - L1 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1671 | 31672 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10SEC | PST (short term flicker)- L2 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1673 | 31674 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10SEC | PST (short term flicker)- L3 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1675 | 31676 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10SEC | PST (short term flicker)- L12 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1677 | 31678 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10SEC | PST (short term flicker)- L23 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1679 | 31680 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10SEC | PST (short term flicker)- L31 - calculated over 10 seconds | 4 | R | 4 Bytes Float | 1681 | 31682 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10MIN | PST (short term flicker)- L1 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1683 | 31684 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10MIN | PST (short term flicker)- L2 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1685 | 31686 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10MIN | PST (short term flicker)- L3 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1687 | 31688 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10MIN | PST (short term flicker)- L12 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1689 | 31690 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10MIN | PST (short term flicker)- L23 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1691 | 31692 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10MIN | PST (short term flicker)- L31 - calculated over 10 minutes | 4 | R | 4 Bytes Float | 1693 | 31694 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1HOUR | PST (short term flicker)- L1 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1695 | 31696 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1HOUR | PLT (Long term flicker)- L2 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1697 | 31698 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1HOUR | PLT (Long term flicker) - L3 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1699 | 31700 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1HOUR | PLT (Long term flicker) - L12 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1701 | 31702 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1HOUR | PLT (Long term flicker) - L23 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1703 | 31704 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1HOUR | PLT (Long term flicker) - L31 - calculated over 1 hour | 4 | R | 4 Bytes Float | 1705 | 31706 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_2HOUR | PLT (Long term flicker)- L1 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1707 | 31708 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_2HOUR | PLT (Long term flicker) - L2 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1709 | 31710 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_2HOUR | PLT (Long term flicker) - L3 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1711 | 31712 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_2HOUR | PLT (Long term flicker) - L12 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1713 | 31714 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_2HOUR | PLT (Long term flicker) - L23 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1715 | 31716 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_2HOUR | PLT (Long term flicker) - L31 - calculated over 2 hours | 4 | R | 4 Bytes Float | 1717 | 31718 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_10HOUR | PLT (Long term flicker) - L1 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1719 | 31720 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_10HOUR | PLT (Long term flicker) - L2 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1721 | 31722 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_10HOUR | PLT (Long term flicker) - L3 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1723 | 31724 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_10HOUR | PLT (Long term flicker) - L12 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1725 | 31726 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_10HOUR | PLT (Long term flicker) - L23 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1727 | 31728 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_10HOUR | PLT (Long term flicker) - L31 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1729 | 31730 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1DAY | PLT (Long term flicker) - L1 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1731 | 31732 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1DAY | PLT (Long term flicker) - L2 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1733 | 31734 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1DAY | PLT (Long term flicker) - L3 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1735 | 31736 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1DAY | PLT (Long term flicker) - L12 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1737 | 31738 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1DAY | PLT (Long term flicker) - L23 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1739 | 31740 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1DAY | PLT (Long term flicker) - L31 - calculated over 10 hours | 4 | R | 4 Bytes Float | 1741 | 31742 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_7DAY | PLT (Long term flicker) - L1 - calculated over 7 days | 4 | R | 4 Bytes Float | 1743 | 31744 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_7DAY | PLT (Long term flicker) - L2 - calculated over 7 days | 4 | R | 4 Bytes Float | 1745 | 31746 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_7DAY | PLT (Long term flicker) - L3 - calculated over 7 days | 4 | R | 4 Bytes Float | 1747 | 31748 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_7DAY | PLT (Long term flicker) - L12 - calculated over 7 days | 4 | R | 4 Bytes Float | 1749 | 31750 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_7DAY | PLT (Long term flicker) - L23 - calculated over 7 days | 4 | R | 4 Bytes Float | 1751 | 31752 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_7DAY | PLT (Long term flicker) - L31 - calculated over 7 days | 4 | R | 4 Bytes Float | 1753 | 31754 | Power Quality Analyzer G4400
PRM_CODE_FLK_TIMESTAMP_10MIN | Flickering's 10 minutes file's begining time stamp (at 100 nsec steps) | 8 | R | 8 Bytes FileTime | 1455 | 31456 | Power Quality Analyzer G4400
PRM_CODE_FLK_TIMESTAMP_2HOUR | Flickering's 2 hours file's beginning time stamp (at 100 nsec steps) | 8 | R | 8 Bytes FileTime | 1459 | 31460 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IN | Total Active energy in since last reset (W) | 8 | R | 8 Bytes Double | 1947 | 31948 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_OUT | Total Active energy out since last reset (W) | 8 | R | 8 Bytes Double | 1951 | 31952 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VA_IN | Total Apparent energy In since last reset (VA) | 8 | R | 8 Bytes Double | 1955 | 31956 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IN | Total Reactive energy out since last reset (VAr) | 8 | R | 8 Bytes Double | 1963 | 31964 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_OUT | Total Reactive energy Out since last reset (VAr) | 8 | R | 8 Bytes Double | 1967 | 31968 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IN |  | 8 | R | 8 Bytes Double | 2740 | 32741 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_OUT |  | 8 | R | 8 Bytes Double | 2744 | 32745 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VA_IN |  | 8 | R | 8 Bytes Double | 2748 | 32749 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IN |  | 8 | R | 8 Bytes Double | 2752 | 32753 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_OUT |  | 8 | R | 8 Bytes Double | 2756 | 32757 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_W_IN | Total Demand Active energy In since last reset (W) | 8 | R | 8 Bytes Double | 1971 | 31972 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_W_OUT | Total Demand Active energy Out since last reset (W) | 8 | R | 8 Bytes Double | 1975 | 31976 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VA_IN | Total Demand Apparent energy In since last reset (VA) | 8 | R | 8 Bytes Double | 1979 | 31980 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VAr_IN | Total Demand Reactive energy In since last reset (VAr) | 8 | R | 8 Bytes Double | 1987 | 31988 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VAr_OUT | Total Demand Reactive energy Out since last reset (VAr) | 8 | R | 8 Bytes Double | 1991 | 31992 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_W_IN | Peak Demand energy in since last reset (W) | 4 | R | 4 Bytes Float | 1995 | 31996 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_W_OUT | Peak demand Active energy out since last reset (W) | 4 | R | 4 Bytes Float | 1997 | 31998 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VA_IN | Peak Demand Apparent energy in since last reset (VA) | 4 | R | 4 Bytes Float | 1999 | 32000 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VAr_IN | Peak Demand Reactive energy in since last reset (VAr) | 4 | R | 4 Bytes Float | 2003 | 32004 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VAr_OUT | Peak Demand Reactive energy out since last reset (VAr) | 4 | R | 4 Bytes Float | 2005 | 32006 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_IN | Accumulated Active energy in of the current period (reset on the beginning of the interval) | 8 | R | 8 Bytes Double | 2007 | 32008 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_OUT | Accumulated Active energy out of the current period (reset on the beginning of the interval) | 8 | R | 8 Bytes Double | 2011 | 32012 | Power Quality Analyzer G4400
PRM_CODE_METERING_VA_IN | Accumulated Apparent energy in of the current period (reset on the beginning of the interval) | 8 | R | 8 Bytes Double | 2015 | 32016 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_IN | Accumulated reactive energy in of the current period (reset on the beginning of the interval) | 8 | R | 8 Bytes Double | 2023 | 32024 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_OUT | Accumulated reactive energy out of the current period (reset on the beginning of the interval) | 8 | R | 8 Bytes Double | 2027 | 32028 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_MEASURE_START | Time stamp start of the energy measurement file (100 nsec steps) | 8 | R | 8 Bytes FileTime | 2031 | 32032 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_MEASURE_LAST_START | Time stamp start of the previous energy measurement interval (100 nsec steps) | 8 | R | 8 Bytes FileTime | 2035 | 32036 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_MEASURE_UP_TIME | Amount of time (delta file time format) the Energy measuring window is so far up (excluding the time the device was off) | 8 | R | 2 Bytes Integer | 2039 | 32040 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_MEASURE_DOWN_TIME | Amount of time (delta file time format) the device was down inside the energy measuring window | 8 | R | 2 Bytes Integer | 2043 | 32044 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_MEASURE_AVAILABILITY | Up time /(up time+down time) | 8 | R | 8 Bytes Double | 2047 | 32048 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IPO | Active energy in (+)Active energy out (wh) | 8 | R | 8 Bytes Double | 2051 | 32052 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IPO | Reactive energy in (+) reactive energy out (VArh) | 8 | R | 8 Bytes Double | 2059 | 32060 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IMO | Active energy in (-) Active energy out (wh) | 8 | R | 8 Bytes Double | 2063 | 32064 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IMO | Reactive energy (-) reactive energy out (VArh) | 8 | R | 8 Bytes Double | 2071 | 32072 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_W_IPO | Demand active energy in (+) demand active energy out (wh) | 8 | R | 8 Bytes Double | 2075 | 32076 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IPO |  | 8 | R | 8 Bytes Double | 2309 | 32310 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IPO |  | 8 | R | 8 Bytes Double | 2313 | 32314 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IMO |  | 8 | R | 8 Bytes Double | 2317 | 32318 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IMO |  | 8 | R | 8 Bytes Double | 2321 | 32322 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VAr_IPO | Demand reactive energy in (+) demand reactive energy out (VArh) | 8 | R | 8 Bytes Double | 2083 | 32084 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_W_IMO | Demand active energy in (-) demand active energy out (wh) | 8 | R | 8 Bytes Double | 2087 | 32088 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VAr_IMO | Demand reactive energy in (-) demand reactive energy out (VArh) | 8 | R | 8 Bytes Double | 2095 | 32096 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_OMI | Accumulated Active energy of the current period (reset on the beginning of the interval) Out minus In | 8 | R | 8 Bytes Double | 5910 | 35911 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_OMI | Total Active energy out (-) Total Active energy in (Wh) | 8 | R | 8 Bytes Double | 5914 | 35915 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_W_OMI | Demand active Energy out (-) Demand active energy in (Wh) | 8 | R | 8 Bytes Double | 5918 | 35919 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_W_OMI | Peak demand active energy out (-) peak demand active energy in (Wh) | 4 | R | 4 Bytes Float | 5920 | 35921 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_OMI | Accumulated reactive energy of the current period (reset on the beginning of the interval) Out minus In | 8 | R | 8 Bytes Double | 5924 | 35925 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_OMI | Total reactive energy out (-) total reactive energy in (Varh) | 8 | R | 8 Bytes Double | 5928 | 35929 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_ENERGY_VAr_OMI | Demand reactive Energy out (-) Demand reactive energy in (Varh) | 8 | R | 8 Bytes Double | 5932 | 35933 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VAr_OMI | Peak demand reactive energy out (-) peak demand reactive energy in (VArh) | 4 | R | 4 Bytes Float | 5936 | 35937 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_OMI_FLP | Total active energy out (-) total active energy in (wh) ==== as 4 bytes float | 4 | R | 4 Bytes Float | 5938 | 35939 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_OMI_FLP | Total reactive energy out (-) total reactive energy in (Varh) ==== as 4 Bytes Float | 4 | R | 4 Bytes Float | 5940 | 35941 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_OUT_FLP |  | 4 | R | 4 Bytes Float | 5942 | 35943 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_OUT_FLP |  | 4 | R | 4 Bytes Float | 5944 | 35945 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_OMI_LAST |  | 4 | R | 4 Bytes Float | 5946 | 35947 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_OMI_LAST |  | 4 | R | 4 Bytes Float | 5948 | 35949 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_IMO_LAST |  | 4 | R | 4 Bytes Float | 5950 | 35951 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_IMO_LAST |  | 4 | R | 4 Bytes Float | 5952 | 35953 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_TIMESTAMP_W_IN |  | 4 | R | 4 Bytes Float | 2696 | 32697 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_TIMESTAMP_W_OUT |  | 4 | R | 4 Bytes Float | 2700 | 32701 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_TIMESTAMP_VA_IN |  | 4 | R | 4 Bytes Float | 2704 | 32705 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_TIMESTAMP_VAr_IN |  | 4 | R | 4 Bytes Float | 2708 | 32709 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_TIMESTAMP_VAr_OUT |  | 4 | R | 4 Bytes Float | 2712 | 32713 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_W_IPO | Peak demand active energy in (+) Peak demand active energy out (Wh) | 4 | R | 4 Bytes Float | 2099 | 32100 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VAr_IPO | Peak demand reactive energy in (+) Peak demand reactive energy out (Varh) | 4 | R | 4 Bytes Float | 2103 | 32104 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_W_IMO | Peak demand active energy in (-) peak demand energy out (Wh) | 4 | R | 4 Bytes Float | 2105 | 32106 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_VAr_IMO | Peak demand reactive energy in (-) Peak demand reactive energy out (Varh) | 4 | R | 4 Bytes Float | 2109 | 32110 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_IPO | Accumulated Active energy of the current period (reset on the beginning of the interval) Input Plus Output | 8 | R | 8 Bytes Double | 2111 | 32112 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_IPO | Accumulated Reactive energy of the current period (reset on the beginning of the interval) Input Plus Output | 8 | R | 8 Bytes Double | 2119 | 32120 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_IMO | Accumulated Active energy of the current period (reset on the beginning of the interval) Input Minus Output | 8 | R | 8 Bytes Double | 2123 | 32124 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_IMO | Accumulated reactive energy of the current period (reset on the beginning of the interval) Input Minus Output | 8 | R | 8 Bytes Double | 2131 | 32132 | Power Quality Analyzer G4400
PRM_CODE_METERING_PF_IN | PF based on metering_in values | 4 | R | 4 Bytes Float | 2135 | 32136 | Power Quality Analyzer G4400
PRM_CODE_METERING_PF_OUT | PF based on metering_out values | 4 | R | 4 Bytes Float | 2137 | 32138 | Power Quality Analyzer G4400
PRM_CODE_METERING_PF_IPO | PF based on metering_in plus out values | 4 | R | 4 Bytes Float | 2139 | 32140 | Power Quality Analyzer G4400
PRM_CODE_METERING_PF_IMO | PF based on metering_in minus out values | 4 | R | 4 Bytes Float | 2141 | 32142 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_PF_IN | Power factor In, calculated every cycle | 4 | R | 4 Bytes Float | 2143 | 32144 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_PF_OUT | Power factor out calculated every cycle | 4 | R | 4 Bytes Float | 2145 | 32146 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_PF_IPO | PF based on metering In plus out | 4 | R | 4 Bytes Float | 2147 | 32148 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_PF_IMO | PF based on metering In Minus out | 4 | R | 4 Bytes Float | 2149 | 32150 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_PF_IN | PF Demand based on metering In | 4 | R | 4 Bytes Float | 2151 | 32152 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_PF_OUT | PF Demand based on metering out | 4 | R | 4 Bytes Float | 2153 | 32154 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_PF_IPO | PF Demand based on metering In plus out | 4 | R | 4 Bytes Float | 2155 | 32156 | Power Quality Analyzer G4400
PRM_CODE_DEMAND_PF_IMO | PF Demand based on metering In minus out | 4 | R | 4 Bytes Float | 2157 | 32158 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_PF_IN | Peak PF Demand based on metering In | 4 | R | 4 Bytes Float | 2159 | 32160 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_PF_OUT | Peak PF Demand based on metering out | 4 | R | 4 Bytes Float | 2161 | 32162 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_PF_IPO | Peak PF Demand based on metering In plus out | 4 | R | 4 Bytes Float | 2163 | 32164 | Power Quality Analyzer G4400
PRM_CODE_PEAK_DEMAND_PF_IMO | Peak PF Demand based on metering In minus out | 4 | R | 4 Bytes Float | 2165 | 32166 | Power Quality Analyzer G4400
PRM_CODE_AVG_DSP_FREQUENCY | Measured frequency average, calculated every 200 msec | 4 | R | 4 Bytes Float | 2191 | 32192 | Power Quality Analyzer G4400
PRM_CODE_AVG_ACTIVE_POWER_MAIN_1 | Active power - average L1 (w), calculated every second | 4 | R | 4 Bytes Float | 2217 | 32218 | Power Quality Analyzer G4400
PRM_CODE_AVG_ACTIVE_POWER_MAIN_2 | Active power - average L2 (w), calculated every second | 4 | R | 4 Bytes Float | 2219 | 32220 | Power Quality Analyzer G4400
PRM_CODE_AVG_ACTIVE_POWER_MAIN_3 | Active power - average L3 (w), calculated every second | 4 | R | 4 Bytes Float | 2221 | 32222 | Power Quality Analyzer G4400
PRM_CODE_AVG_ACTIVE_POWER_SUM | Active power - total (w), calculated every second | 4 | R | 4 Bytes Float | 2229 | 32230 | Power Quality Analyzer G4400
PRM_CODE_AVG_REACTIVE_POWER_MAIN_1 | Reactive power - average L1 (VAr), calculated every second | 4 | R | 4 Bytes Float | 2233 | 32234 | Power Quality Analyzer G4400
PRM_CODE_AVG_REACTIVE_POWER_MAIN_2 | Reactive power - average L2 (VAr), calculated every second | 4 | R | 4 Bytes Float | 2235 | 32236 | Power Quality Analyzer G4400
PRM_CODE_AVG_REACTIVE_POWER_MAIN_3 | Reactive power - average L3 (VAr), calculated every second | 4 | R | 4 Bytes Float | 2237 | 32238 | Power Quality Analyzer G4400
PRM_CODE_AVG_REACTIVE_POWER_SUM | Reactive power - total (VAr), calculated every second | 4 | R | 4 Bytes Float | 2245 | 32246 | Power Quality Analyzer G4400
PRM_CODE_AVG_V1_RMS | RMS Voltage V1 - average (v) calculated every second | 4 | R | 4 Bytes Float | 2283 | 32284 | Power Quality Analyzer G4400
PRM_CODE_AVG_V2_RMS | RMS Voltage V2 - average (v), calculated every second | 4 | R | 4 Bytes Float | 2285 | 32286 | Power Quality Analyzer G4400
PRM_CODE_AVG_V3_RMS | RMS Voltage V3 - average (v), calculated every second | 4 | R | 4 Bytes Float | 2287 | 32288 | Power Quality Analyzer G4400
PRM_CODE_AVG_V12_RMS | RMS Voltage V12 - average (v), calculated every second | 4 | R | 4 Bytes Float | 2289 | 32290 | Power Quality Analyzer G4400
PRM_CODE_AVG_V23_RMS | RMS Voltage V23 - average (v), calculated every second | 4 | R | 4 Bytes Float | 2291 | 32292 | Power Quality Analyzer G4400
PRM_CODE_AVG_V31_RMS | RMS Voltage V31 - average (v), calculated every second | 4 | R | 4 Bytes Float | 2293 | 32294 | Power Quality Analyzer G4400
PRM_CODE_AVG_I1_RMS | RMS Voltage I1 - average (A), calculated every second | 4 | R | 4 Bytes Float | 2297 | 32298 | Power Quality Analyzer G4400
PRM_CODE_AVG_I2_RMS | RMS Voltage I2 - average (A), calculated every second | 4 | R | 4 Bytes Float | 2299 | 32300 | Power Quality Analyzer G4400
PRM_CODE_AVG_I3_RMS | RMS current I3 - average (A), calculated every second | 4 | R | 4 Bytes Float | 2301 | 32302 | Power Quality Analyzer G4400
PRM_CODE_AVG_V1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Voltage | 4 | R | 4 Bytes Float | 2303 | 32304 | Power Quality Analyzer G4400
PRM_CODE_AVG_V12_23_31_RMS_AVERAGE | Cycle by cycle Average RMS phase to phase Voltage | 4 | R | 4 Bytes Float | 2305 | 32306 | Power Quality Analyzer G4400
PRM_CODE_AVG_I1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Current | 4 | R | 4 Bytes Float | 2307 | 32308 | Power Quality Analyzer G4400
PRM_CODE_AVG_TEMPERATURE_DSP_INTERNAL | Internal Average Temperature at the DSP module, averaged over 1 second intervals since last reset | 4 | R | 4 Bytes Float | 2339 | 32340 | Power Quality Analyzer G4400
PRM_CODE_AVG_TEMPERATURE_DSP_EXTERNAL | External Average Temperature at the DSP module (PT100), averaged over 1 second intervals since last reset | 4 | R | 4 Bytes Float | 2341 | 32342 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_DSP_FREQUENCY | Actual frequency as part of the Nominal frequency (%) calculated every 200 msec | 4 | R | 4 Bytes Float | 2343 | 32344 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_ACTIVE_POWER_N | Active power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2345 | 32346 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_ACTIVE_POWER_MAIN_1 | Active power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2347 | 32348 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_ACTIVE_POWER_MAIN_2 | Active power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2349 | 32350 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_ACTIVE_POWER_MAIN_3 | Active power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2351 | 32352 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_ACTIVE_POWER_SUM | Total active power as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2359 | 32360 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_REACTIVE_POWER_N | Average reactive power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2361 | 32362 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_REACTIVE_POWER_MAIN_1 | Average Reactive power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2363 | 32364 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_REACTIVE_POWER_MAIN_2 | Average Reactive power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2365 | 32366 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_REACTIVE_POWER_MAIN_3 | Average Reactive power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2367 | 32368 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_REACTIVE_POWER_SUM | Total Average reactive power as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2375 | 32376 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_APPARENT_POWER_N | Average Apparent power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2377 | 32378 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_APPARENT_POWER_MAIN_1 | Average Apparent power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2379 | 32380 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_APPARENT_POWER_MAIN_2 | Average Apparent power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2381 | 32382 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_APPARENT_POWER_MAIN_3 | Average Apparent power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2383 | 32384 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_APPARENT_POWER_SUM | Total Average Apparent power as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 2391 | 32392 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_POWER_FACTOR_N |  | 4 | R | 4 Bytes Float | 2393 | 32394 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_POWER_FACTOR_1 |  | 4 | R | 4 Bytes Float | 2395 | 32396 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_POWER_FACTOR_2 |  | 4 | R | 4 Bytes Float | 2397 | 32398 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_POWER_FACTOR_3 |  | 4 | R | 4 Bytes Float | 2399 | 32400 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_POWER_FACTOR_TOTAL |  | 4 | R | 4 Bytes Float | 2401 | 32402 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_PF_FUNDAMENTAL_1 |  | 4 | R | 4 Bytes Float | 2403 | 32404 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_PF_FUNDAMENTAL_2 |  | 4 | R | 4 Bytes Float | 2405 | 32406 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_PF_FUNDAMENTAL_3 |  | 4 | R | 4 Bytes Float | 2407 | 32408 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_PF_FUNDAMENTAL_TOTAL |  | 4 | R | 4 Bytes Float | 2409 | 32410 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_VN_RMS | VN RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2411 | 32412 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V1_RMS | V1 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2413 | 32414 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V2_RMS | V2 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2415 | 32416 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V3_RMS | V3 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2417 | 32418 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V12_RMS | V12 RMS voltage averaged over 1 second intervals as part of it's Nominal (%) | 4 | R | 4 Bytes Float | 2419 | 32420 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V23_RMS | V23 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2421 | 32422 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V31_RMS | V31 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2423 | 32424 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_IN_RMS | IN RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2425 | 32426 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_I1_RMS | I1 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2427 | 32428 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_I2_RMS | I2 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2429 | 32430 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_I3_RMS | I3 RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2431 | 32432 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Voltage as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2433 | 32434 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_V12_23_31_RMS_AVERAGE | Cycle by cycle Average RMS phase to phase Voltage as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2435 | 32436 | Power Quality Analyzer G4400
PRM_CODE_AVG_PU_I1_2_3_RMS_AVERAGE | Cycle by cycle Average RMS Current as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 2437 | 32438 | Power Quality Analyzer G4400
PRM_CODE_SAVED_INTERHARMONIC_PRM |  | 2 | RW | 2 Bytes Integer | 2439 | 32440 | Power Quality Analyzer G4400
PRM_CODE_SAVED_INTERHARMONIC_INDEX |  | 2 | RW | 2 Bytes Integer | 2440 | 32441 | Power Quality Analyzer G4400
PRM_CODE_SAVED_HARMONICS_PRM |  | 2 | RW | 2 Bytes Integer | 2441 | 32442 | Power Quality Analyzer G4400
PRM_CODE_HARMONICS_VALUE |  | 200 | R | 4 Bytes Float | 2442 | 32443 | Power Quality Analyzer G4400
PRM_CODE_INTERHARMONIC_VALUE |  | 4 | R | 4 Bytes Float | 2542 | 32543 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_2SEC_MIN | V1 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1755 | 31756 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_2SEC_MIN | V2 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1757 | 31758 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_2SEC_MIN | V3 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1759 | 31760 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_2SEC_MIN | V12 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1761 | 31762 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_2SEC_MIN | V23 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1763 | 31764 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_2SEC_MIN | V31 Min value of the Short term flicker, calculated over 2 seconds | 4 | R | 4 Bytes Float | 1765 | 31766 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10SEC_MIN | V1 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1767 | 31768 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10SEC_MIN | V2 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1769 | 31770 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10SEC_MIN | V3 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1771 | 31772 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10SEC_MIN | V12 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1773 | 31774 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10SEC_MIN | V23 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1775 | 31776 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10SEC_MIN | V31 Min value of the Short term flicker, calculated over 10 seconds | 4 | R | 4 Bytes Float | 1777 | 31778 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10MIN_MIN | V1 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1779 | 31780 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10MIN_MIN | V2 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1781 | 31782 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10MIN_MIN | V3 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1783 | 31784 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10MIN_MIN | V12 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1785 | 31786 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10MIN_MIN | V23 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1787 | 31788 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10MIN_MIN | V31 Min value of the Short term flicker, calculated over 10 minutes | 4 | R | 4 Bytes Float | 1789 | 31790 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1HOUR_MIN | V1 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1791 | 31792 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1HOUR_MIN | V2 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1793 | 31794 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1HOUR_MIN | V3 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1795 | 31796 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1HOUR_MIN | V12 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1797 | 31798 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1HOUR_MIN | V23 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1799 | 31800 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1HOUR_MIN | V31 Min value of the long term flicker, calculated over 1 hour | 4 | R | 4 Bytes Float | 1801 | 31802 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_2HOUR_MIN | V1 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1803 | 31804 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_2HOUR_MIN | V2 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1805 | 31806 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_2HOUR_MIN | V3 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1807 | 31808 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_2HOUR_MIN | V12 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1809 | 31810 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_2HOUR_MIN | V23 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1811 | 31812 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_2HOUR_MIN | V31 Min value of the long term flicker, calculated over 2 hour | 4 | R | 4 Bytes Float | 1813 | 31814 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_10HOUR_MIN | V1 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1815 | 31816 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_10HOUR_MIN | V2 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1817 | 31818 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_10HOUR_MIN | V3 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1819 | 31820 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_10HOUR_MIN | V12 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1821 | 31822 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_10HOUR_MIN | V23 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1823 | 31824 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_10HOUR_MIN | V31 Min value of the long term flicker, calculated over 10 hour | 4 | R | 4 Bytes Float | 1825 | 31826 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1DAY_MIN | V1 Min value of the long term flicker, calculated over 1 day | 4 | R | 4 Bytes Float | 1827 | 31828 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1DAY_MIN | V2 Min value of the long term flicker, calculated over 1 day | 4 | R | 4 Bytes Float | 1829 | 31830 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1DAY_MIN | V3 Min value of the long term flicker, calculated over 1 day | 4 | R | 4 Bytes Float | 1831 | 31832 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1DAY_MIN | V12 Min value of the long term flicker, calculated over 1day | 4 | R | 4 Bytes Float | 1833 | 31834 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1DAY_MIN | V23 Min value of the long term flicker, calculated over 1 day | 4 | R | 4 Bytes Float | 1835 | 31836 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1DAY_MIN | V31 Min value of the long term flicker, calculated over 1 day | 4 | R | 4 Bytes Float | 1837 | 31838 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_7DAY_MIN | V1 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1839 | 31840 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_7DAY_MIN | V2 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1841 | 31842 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_7DAY_MIN | V3 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1843 | 31844 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_7DAY_MIN | V12 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1845 | 31846 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_7DAY_MIN | V23 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1847 | 31848 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_7DAY_MIN | V31 Min value of the long term flicker, calculated over 7 days | 4 | R | 4 Bytes Float | 1849 | 31850 | Power Quality Analyzer G4400
PRM_CODE_PSU_TEMPERATURE_MIN | Minimum reading (since last reset) of the Power supply temperature, calculated every cycle | 4 | R | 4 Bytes Float | 539 | 30540 | Power Quality Analyzer G4400
PRM_CODE_DSP_FREQUENCY_MIN | Minimum measured frequency, calculated every 200 msec | 4 | R | 4 Bytes Float | 1163 | 31164 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_1_MIN | Active power - Minimum reading L1 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1189 | 31190 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_2_MIN | Active power - Minimum reading L2 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1191 | 31192 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_3_MIN | Active power - Minimum reading L3 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1193 | 31194 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_SUM_MIN | Active power - Minimum total reading (w) calculated every cycle | 4 | R | 4 Bytes Float | 1201 | 31202 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_1_MIN | Minimum Reactive power - L1 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1205 | 31206 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_2_MIN | Minimum Reactive power - L2 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1207 | 31208 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_3_MIN | Minimum Reactive power - L3 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1209 | 31210 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_SUM_MIN | Minimum Reactive power - total (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1217 | 31218 | Power Quality Analyzer G4400
PRM_CODE_V1_RMS_MIN | Minimum V1 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1255 | 31256 | Power Quality Analyzer G4400
PRM_CODE_V2_RMS_MIN | Minimum V2 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1257 | 31258 | Power Quality Analyzer G4400
PRM_CODE_V3_RMS_MIN | Minimum V3 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1259 | 31260 | Power Quality Analyzer G4400
PRM_CODE_V12_RMS_MIN | Minimum V12 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1261 | 31262 | Power Quality Analyzer G4400
PRM_CODE_V23_RMS_MIN | Minimum V23 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1263 | 31264 | Power Quality Analyzer G4400
PRM_CODE_V31_RMS_MIN | Minimum V31 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1265 | 31266 | Power Quality Analyzer G4400
PRM_CODE_I1_RMS_MIN | Minimum I1 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1269 | 31270 | Power Quality Analyzer G4400
PRM_CODE_I2_RMS_MIN | Minimum I2 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1271 | 31272 | Power Quality Analyzer G4400
PRM_CODE_I3_RMS_MIN | Minimum I3 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1273 | 31274 | Power Quality Analyzer G4400
PRM_CODE_V1_2_3_RMS_AVERAGE_MIN | Cycle by cycle Minimum RMS Voltage | 4 | R | 4 Bytes Float | 1275 | 31276 | Power Quality Analyzer G4400
PRM_CODE_V12_23_31_RMS_AVERAGE_MIN | Cycle by cycle Minimum RMS phase to phase Voltage | 4 | R | 4 Bytes Float | 1277 | 31278 | Power Quality Analyzer G4400
PRM_CODE_I1_2_3_RMS_AVERAGE_MIN | Cycle by cycle Minimum RMS Current | 4 | R | 4 Bytes Float | 1279 | 31280 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_INTERNAL_MIN | Internal Temperature at the DSP module , calculated over 1 cycle, minimum value since last reset | 4 | R | 4 Bytes Float | 1311 | 31312 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_EXTERNAL_MIN | external Temperature at the DSP module , calculated over 1 cycle, minimum value since last reset | 4 | R | 4 Bytes Float | 1313 | 31314 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_2SEC_MAX | V1 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1851 | 31852 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_2SEC_MAX | V2 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1853 | 31854 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_2SEC_MAX | V3 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1855 | 31856 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_2SEC_MAX | V12 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1857 | 31858 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_2SEC_MAX | V23 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1859 | 31860 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_2SEC_MAX | V31 Short term flicker Max value (since last reset) calculated over 2 seconds | 4 | R | 4 Bytes Float | 1861 | 31862 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10SEC_MAX | V1 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1863 | 31864 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10SEC_MAX | V2 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1865 | 31866 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10SEC_MAX | V3 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1867 | 31868 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10SEC_MAX | V12 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1869 | 31870 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10SEC_MAX | V23 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1871 | 31872 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10SEC_MAX | V31 Short term flicker Max value (since last reset) calculated over 10 seconds | 4 | R | 4 Bytes Float | 1873 | 31874 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PST_10MIN_MAX | V1 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1875 | 31876 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PST_10MIN_MAX | V2 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1877 | 31878 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PST_10MIN_MAX | V3 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1879 | 31880 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PST_10MIN_MAX | V12 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1881 | 31882 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PST_10MIN_MAX | V23 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1883 | 31884 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PST_10MIN_MAX | V31 Short term flicker Max value (since last reset) calculated over 10 minutes | 4 | R | 4 Bytes Float | 1885 | 31886 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1HOUR_MAX | V1 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1887 | 31888 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1HOUR_MAX | V2 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1889 | 31890 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1HOUR_MAX | V3 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1891 | 31892 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1HOUR_MAX | V12 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1893 | 31894 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1HOUR_MAX | V23 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1895 | 31896 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1HOUR_MAX | V31 long term flicker Max value (since last reset) calculated over 1 hour | 4 | R | 4 Bytes Float | 1897 | 31898 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_2HOUR_MAX | V1 long term flicker Max value (since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1899 | 31900 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_2HOUR_MAX | V2 long term flicker Max value (since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1901 | 31902 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_2HOUR_MAX | V3 long term flicker Max value (since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1903 | 31904 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_2HOUR_MAX | V12 long term flicker Max value (since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1905 | 31906 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_2HOUR_MAX | V23 long term flicker Max value (since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1907 | 31908 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_2HOUR_MAX | V31 long term flicker Max value(since last reset) calculated over 2 hour | 4 | R | 4 Bytes Float | 1909 | 31910 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_10HOUR_MAX | V1 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1911 | 31912 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_10HOUR_MAX | V2 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1913 | 31914 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_10HOUR_MAX | V3 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1915 | 31916 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_10HOUR_MAX | V12 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1917 | 31918 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_10HOUR_MAX | V23 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1919 | 31920 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_10HOUR_MAX | V31 long term flicker Max value (since last reset) calculated over 10 hour | 4 | R | 4 Bytes Float | 1921 | 31922 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_1DAY_MAX | V1 long term flicker Max value (since last reset) calculated over 1 day | 4 | R | 4 Bytes Float | 1923 | 31924 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_1DAY_MAX | V2 long term flicker Max value (since last reset) calculated over 1 day | 4 | R | 4 Bytes Float | 1925 | 31926 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_1DAY_MAX | V3 long term flicker Max value (since last reset) calculated over 1 day | 4 | R | 4 Bytes Float | 1927 | 31928 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_1DAY_MAX | V12 long term flicker Max value (since last reset) calculated over 1day | 4 | R | 4 Bytes Float | 1929 | 31930 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_1DAY_MAX | V23 long term flicker Max value (since last reset) calculated over 1 day | 4 | R | 4 Bytes Float | 1931 | 31932 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_1DAY_MAX | V31 long term flicker Max value (since last reset) calculated over 1 day | 4 | R | 4 Bytes Float | 1933 | 31934 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V1_PLT_7DAY_MAX | V1 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1935 | 31936 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V2_PLT_7DAY_MAX | V2 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1937 | 31938 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V3_PLT_7DAY_MAX | V3 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1939 | 31940 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V12_PLT_7DAY_MAX | V12 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1941 | 31942 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V23_PLT_7DAY_MAX | V23 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1943 | 31944 | Power Quality Analyzer G4400
PRM_CODE_FLICKERING_V31_PLT_7DAY_MAX | V31 long term flicker Max value (since last reset) calculated over 7 days | 4 | R | 4 Bytes Float | 1945 | 31946 | Power Quality Analyzer G4400
PRM_CODE_PSU_TEMPERATURE_MAX | Power supply temperaturt, calculated over 1 cycle, maximum value since last reset | 4 | R | 4 Bytes Float | 541 | 30542 | Power Quality Analyzer G4400
PRM_CODE_DSP_FREQUENCY_MAX | Maximum frequency measured, calculated every 200 msec | 4 | R | 4 Bytes Float | 1315 | 31316 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_1_MAX | Active power - Maximum reading L1 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1341 | 31342 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_2_MAX | Active power - Maximum reading L2 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1343 | 31344 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_POWER_MAIN_3_MAX | Active power - Maximum reading L3 (w) calculated every cycle | 4 | R | 4 Bytes Float | 1345 | 31346 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_1_MAX | Maximum Reactive power - L1 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1357 | 31358 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_2_MAX | Maximum Reactive power - L2 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1359 | 31360 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_MAIN_3_MAX | Maximum Reactive power - L3 (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1361 | 31362 | Power Quality Analyzer G4400
PRM_CODE_REACTIVE_POWER_SUM_MAX | Maximum Reactive power - total (VAr) calculated every cycle | 4 | R | 4 Bytes Float | 1369 | 31370 | Power Quality Analyzer G4400
PRM_CODE_V1_RMS_MAX | Maximum V1 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1407 | 31408 | Power Quality Analyzer G4400
PRM_CODE_V2_RMS_MAX | Maximum V2 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1409 | 31410 | Power Quality Analyzer G4400
PRM_CODE_V3_RMS_MAX | Maximum V3 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1411 | 31412 | Power Quality Analyzer G4400
PRM_CODE_V12_RMS_MAX | Maximum V12 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1413 | 31414 | Power Quality Analyzer G4400
PRM_CODE_V23_RMS_MAX | Maximum V23 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1415 | 31416 | Power Quality Analyzer G4400
PRM_CODE_V31_RMS_MAX | Maximum V31 RMS reading, calculated every cycle (v) | 4 | R | 4 Bytes Float | 1417 | 31418 | Power Quality Analyzer G4400
PRM_CODE_I1_RMS_MAX | Maximum I1 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1421 | 31422 | Power Quality Analyzer G4400
PRM_CODE_I2_RMS_MAX | Maximum I2 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1423 | 31424 | Power Quality Analyzer G4400
PRM_CODE_I3_RMS_MAX | Maximum I3 RMS reading, calculated every cycle (A) | 4 | R | 4 Bytes Float | 1425 | 31426 | Power Quality Analyzer G4400
PRM_CODE_V1_2_3_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS Voltage | 4 | R | 4 Bytes Float | 1427 | 31428 | Power Quality Analyzer G4400
PRM_CODE_V12_23_31_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS phase to phase Voltage | 4 | R | 4 Bytes Float | 1429 | 31430 | Power Quality Analyzer G4400
PRM_CODE_I1_2_3_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS Current | 4 | R | 4 Bytes Float | 1431 | 31432 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_INTERNAL_MAX | DSP internal temperaturt, calculated over 1 cycle, maximum measurement since last reset | 4 | R | 4 Bytes Float | 1463 | 31464 | Power Quality Analyzer G4400
PRM_CODE_TEMPERATURE_DSP_EXTERNAL_MAX | DSP external temperaturt, calculated over 1 cycle, maximum measurement since last reset | 4 | R | 4 Bytes Float | 1465 | 31466 | Power Quality Analyzer G4400
PRM_CODE_PU_DSP_FREQUENCY_MIN | Minimum frequency measured as part from the nominal (%) | 4 | R | 4 Bytes Float | 1467 | 31468 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_N_MIN | Minimum Active power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1469 | 31470 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_1_MIN | Minimum Active power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1471 | 31472 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_2_MIN | Minimum Active power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1473 | 31474 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_3_MIN | Minimum Active power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1475 | 31476 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_SUM_MIN | Minimum active power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1483 | 31484 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_N_MIN | Minimum reactive power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1485 | 31486 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_1_MIN | Minimum reactive power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1487 | 31488 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_2_MIN | Minimum reactive power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1489 | 31490 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_3_MIN | Minimum reactive power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1491 | 31492 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_SUM_MIN | Minimum reactive power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1499 | 31500 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_N_MIN | Minimum apparent power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1501 | 31502 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_1_MIN | Minimum apparent power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1503 | 31504 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_2_MIN | Minimum apparent power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1505 | 31506 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_3_MIN | Minimum apparent power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1507 | 31508 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_SUM_MIN | Minimum apparent active power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1515 | 31516 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_N_MIN |  | 4 | R | 4 Bytes Float | 1517 | 31518 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_N_MIN |  | 4 | R | 4 Bytes Float | 1517 | 31518 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_1_MIN |  | 4 | R | 4 Bytes Float | 1519 | 31520 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_2_MIN |  | 4 | R | 4 Bytes Float | 1521 | 31522 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_3_MIN |  | 4 | R | 4 Bytes Float | 1523 | 31524 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_TOTAL_MIN |  | 4 | R | 4 Bytes Float | 1525 | 31526 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_1_MIN |  | 4 | R | 4 Bytes Float | 1527 | 31528 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_2_MIN |  | 4 | R | 4 Bytes Float | 1529 | 31530 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_3_MIN |  | 4 | R | 4 Bytes Float | 1531 | 31532 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_TOTAL_MIN |  | 4 | R | 4 Bytes Float | 1533 | 31534 | Power Quality Analyzer G4400
PRM_CODE_PU_VN_RMS_MIN | VN Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1535 | 31536 | Power Quality Analyzer G4400
PRM_CODE_PU_V1_RMS_MIN | V1 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1537 | 31538 | Power Quality Analyzer G4400
PRM_CODE_PU_V2_RMS_MIN | V2 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1539 | 31540 | Power Quality Analyzer G4400
PRM_CODE_PU_V3_RMS_MIN | V3 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1541 | 31542 | Power Quality Analyzer G4400
PRM_CODE_PU_V12_RMS_MIN | V12 Minimum RMS voltage averaged over 1 second intervals as part of it's Nominal (%) | 4 | R | 4 Bytes Float | 1543 | 31544 | Power Quality Analyzer G4400
PRM_CODE_PU_V23_RMS_MIN | V23 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1545 | 31546 | Power Quality Analyzer G4400
PRM_CODE_PU_V31_RMS_MIN | V31 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1547 | 31548 | Power Quality Analyzer G4400
PRM_CODE_PU_IN_RMS_MIN | IN Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1549 | 31550 | Power Quality Analyzer G4400
PRM_CODE_PU_I1_RMS_MIN | I1 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1551 | 31552 | Power Quality Analyzer G4400
PRM_CODE_PU_I2_RMS_MIN | I2 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1553 | 31554 | Power Quality Analyzer G4400
PRM_CODE_PU_I3_RMS_MIN | I3 Minimum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1555 | 31556 | Power Quality Analyzer G4400
PRM_CODE_PU_V1_2_3_RMS_AVERAGE_MIN | Minimum of the average of V1-V2-V3 as part of the nominal since last reset | 4 | R | 4 Bytes Float | 1557 | 31558 | Power Quality Analyzer G4400
PRM_CODE_PU_V12_23_31_RMS_AVERAGE_MIN | Minimum of the average of V12-V23-V31 as part of the nominal since last reset | 4 | R | 4 Bytes Float | 1559 | 31560 | Power Quality Analyzer G4400
PRM_CODE_PU_I1_2_3_RMS_AVERAGE_MIN | Minimum of the average of I1-I2-I3 as part of the nominal since last reset | 4 | R | 4 Bytes Float | 1561 | 31562 | Power Quality Analyzer G4400
PRM_CODE_PU_DSP_FREQUENCY_MAX | Network maximum measurement over 1 cycle as part of the nominal (%) | 4 | R | 4 Bytes Float | 1563 | 31564 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_N_MAX | Maximum Active power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1565 | 31566 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_1_MAX | Maximum Active power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1567 | 31568 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_2_MAX | Maximum Active power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1569 | 31570 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_MAIN_3_MAX | Maximum Active power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1571 | 31572 | Power Quality Analyzer G4400
PRM_CODE_PU_ACTIVE_POWER_SUM_MAX | Maximum active power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1579 | 31580 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_N_MAX | Maximum reactive power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1581 | 31582 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_1_MAX | Maximum reactive power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1583 | 31584 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_2_MAX | Maximum reactive power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1585 | 31586 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_MAIN_3_MAX | Maximum reactive power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1587 | 31588 | Power Quality Analyzer G4400
PRM_CODE_PU_REACTIVE_POWER_SUM_MAX | Maximum reactive power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1595 | 31596 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_N_MAX | Maximum apparent power N reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1597 | 31598 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_1_MAX | Maximum apparent power phase 1 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1599 | 31600 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_2_MAX | Maximum apparent power phase 2 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1601 | 31602 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_MAIN_3_MAX | Maximum apparent power phase 3 reading as part of It's nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1603 | 31604 | Power Quality Analyzer G4400
PRM_CODE_PU_APPARENT_POWER_SUM_MAX | Maximum apparent active power (total) as part of the nominal (%) calculated every second | 4 | R | 4 Bytes Float | 1611 | 31612 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_N_MAX |  | 4 | R | 4 Bytes Float | 1613 | 31614 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_1_MAX |  | 4 | R | 4 Bytes Float | 1615 | 31616 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_2_MAX |  | 4 | R | 4 Bytes Float | 1617 | 31618 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_3_MAX |  | 4 | R | 4 Bytes Float | 1619 | 31620 | Power Quality Analyzer G4400
PRM_CODE_PU_POWER_FACTOR_TOTAL_MAX |  | 4 | R | 4 Bytes Float | 1621 | 31622 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_1_MAX |  | 4 | R | 4 Bytes Float | 1623 | 31624 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_2_MAX |  | 4 | R | 4 Bytes Float | 1625 | 31626 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_3_MAX |  | 4 | R | 4 Bytes Float | 1627 | 31628 | Power Quality Analyzer G4400
PRM_CODE_PU_PF_FUNDAMENTAL_TOTAL_MAX |  | 4 | R | 4 Bytes Float | 1629 | 31630 | Power Quality Analyzer G4400
PRM_CODE_PU_VN_RMS_MAX | VN Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1631 | 31632 | Power Quality Analyzer G4400
PRM_CODE_PU_V1_RMS_MAX | V1 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1633 | 31634 | Power Quality Analyzer G4400
PRM_CODE_PU_V2_RMS_MAX | V2 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1635 | 31636 | Power Quality Analyzer G4400
PRM_CODE_PU_V3_RMS_MAX | V3 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1637 | 31638 | Power Quality Analyzer G4400
PRM_CODE_PU_V12_RMS_MAX | V12 Maximum RMS voltage averaged over 1 second intervals as part of it's Nominal (%) | 4 | R | 4 Bytes Float | 1639 | 31640 | Power Quality Analyzer G4400
PRM_CODE_PU_V23_RMS_MAX | V23 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1641 | 31642 | Power Quality Analyzer G4400
PRM_CODE_PU_V31_RMS_MAX | V31 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1643 | 31644 | Power Quality Analyzer G4400
PRM_CODE_PU_IN_RMS_MAX | IN Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1645 | 31646 | Power Quality Analyzer G4400
PRM_CODE_PU_I1_RMS_MAX | I1 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1647 | 31648 | Power Quality Analyzer G4400
PRM_CODE_PU_I2_RMS_MAX | I2 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1649 | 31650 | Power Quality Analyzer G4400
PRM_CODE_PU_I3_RMS_MAX | I3 Maximum RMS voltage averaged over 1 second intervals as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1651 | 31652 | Power Quality Analyzer G4400
PRM_CODE_PU_V1_2_3_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS Voltage as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1653 | 31654 | Power Quality Analyzer G4400
PRM_CODE_PU_V12_23_31_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS phase to phase Voltage as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1655 | 31656 | Power Quality Analyzer G4400
PRM_CODE_PU_I1_2_3_RMS_AVERAGE_MAX | Cycle by cycle Maximum RMS Current as part of It's Nominal (%) | 4 | R | 4 Bytes Float | 1657 | 31658 | Power Quality Analyzer G4400
PRM_CODE_PHASE_ORDER | Voltage connections sequence | 4 | R | 4 Bytes Integer | 3399 | 33400 | Power Quality Analyzer G4400
PRM_CODE_PHASORS4MODUBS |  | 0 | R | UNKNOWN!!!! | 2640 | 32641 | Power Quality Analyzer G4400
PRM_CODE_PRESET_DATA_BLOCK |  | 20 | R | UNKNOWN!!!! | 1219 | 31220 | Power Quality Analyzer G4400
PRM_CODE_RESET_MINMAX_MODBUS |  | 0 | RW | Action | 5211 | 35212 | Power Quality Analyzer G4400
PRM_CODE_RESET_ENERGY_MODBUS |  | 0 | RW | Action | 5212 | 35213 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10SEC_FREQUENCY | Frequency over 10 seconds, calculated from 10/12 cycles intervals | 4 | R | 4 Bytes Float | 5000 | 35001 | Power Quality Analyzer G4400
PRM_CODE_AVG_TIMESTAMP_10SEC | Start Time of the 10 second interval (at a 100 nsec steps accuracy) | 8 | R | 8 Bytes FileTime | 5002 | 35003 | Power Quality Analyzer G4400
PRM_CODE_AVG_FLAG_2HOURS | Indicate an trigger for event, within the last checked 2 hours time interval | 4 | R | 4 Bytes Integer | 5006 | 35007 | Power Quality Analyzer G4400
PRM_CODE_AVG_TIMESTAMP_2HOURS | Start Time of the 2 houre interval (at a 100 nsec steps accuracy) | 8 | R | 8 Bytes FileTime | 5008 | 35009 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_VN_RMS | VN RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5012 | 35013 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V1_RMS | V1 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5014 | 35015 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V2_RMS | V2 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5016 | 35017 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V3_RMS | V3 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5018 | 35019 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V12_RMS | V12 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5020 | 35021 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V23_RMS | V23 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5022 | 35023 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V31_RMS | V31 RMS aggregated value over 2 hours, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5024 | 35025 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V_UNBALANCE_ZERO_RATIO | 2 Hours averaging - u0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5026 | 35027 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_I_UNBALANCE_ZERO_RATIO | 2 Hours averaging - i0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5028 | 35029 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V_UNBALANCE | V average unbalance calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5030 | 35031 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_I_UNBALANCE | I average unbalance calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5032 | 35033 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V_POSITIVE_UNBALANCE | V average positive sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5034 | 35035 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_I_POSITIVE_UNBALANCE | I average positive sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5036 | 35037 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V_NEGATIVE_UNBALANCE | V average negative sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5038 | 35039 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_I_NEGATIVE_UNBALANCE | I average negative sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5040 | 35041 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V_ZERO_UNBALANCE | V average zero sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5042 | 35043 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_I_ZERO_UNBALANCE | I average zero sequence calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5044 | 35045 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_VN_OVERDEVIATION | VN average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5046 | 35047 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V1_OVERDEVIATION | V1 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5048 | 35049 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V2_OVERDEVIATION | V2 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5050 | 35051 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V3_OVERDEVIATION | V3 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5052 | 35053 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V12_OVERDEVIATION | V12 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5054 | 35055 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V23_OVERDEVIATION | V23 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5056 | 35057 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V31_OVERDEVIATION | V31 average over deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5058 | 35059 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_VN_UNDERDEVIATION | VN average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5060 | 35061 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V1_UNDERDEVIATION | V1 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5062 | 35063 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V2_UNDERDEVIATION | V2 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5064 | 35065 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V3_UNDERDEVIATION | V3 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5066 | 35067 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V12_UNDERDEVIATION | V12 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5068 | 35069 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V23_UNDERDEVIATION | V23 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5070 | 35071 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOURS_V31_UNDERDEVIATION | V31 average under deviation calculated over 2 hours by 10/12 cycles values | 4 | R | 4 Bytes Float | 5072 | 35073 | Power Quality Analyzer G4400
PRM_CODE_AVG_FLAG_3SEC | Indicate an trigger for event, within the last checked 3 seconds time interval | 4 | R | 4 Bytes Integer | 5074 | 35075 | Power Quality Analyzer G4400
PRM_CODE_AVG_TIMESTAMP_3SEC | Start Time of the 3 seconds interval (at a 100 nsec steps accuracy) | 8 | R | 8 Bytes FileTime | 5076 | 35077 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_VN_RMS | VN RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5080 | 35081 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V1_RMS | V1 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5082 | 35083 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V2_RMS | V2 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5084 | 35085 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V3_RMS | V3 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5086 | 35087 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V12_RMS | V12 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5088 | 35089 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V23_RMS | V23 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5090 | 35091 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V31_RMS | V31 RMS aggregated value over 3 seconds, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5092 | 35093 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V_UNBALANCE_ZERO_RATIO | 150/180 cycles averaging - u0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5094 | 35095 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_I_UNBALANCE_ZERO_RATIO | 150/180 cycles averaging - i0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5096 | 35097 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V_UNBALANCE | V average unbalance calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5098 | 35099 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_I_UNBALANCE | I average unbalance calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5100 | 35101 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V_POSITIVE_UNBALANCE | V average positive sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5102 | 35103 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_I_POSITIVE_UNBALANCE | I average positive sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5104 | 35105 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V_NEGATIVE_UNBALANCE | V average negative sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5106 | 35107 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_I_NEGATIVE_UNBALANCE | I average negative sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5108 | 35109 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V_ZERO_UNBALANCE | V average zero sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5110 | 35111 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_I_ZERO_UNBALANCE | I average zero sequence calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5112 | 35113 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_VN_OVERDEVIATION | VN average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5114 | 35115 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V1_OVERDEVIATION | V1 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5116 | 35117 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V2_OVERDEVIATION | V2 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5118 | 35119 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V3_OVERDEVIATION | V3 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5120 | 35121 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V12_OVERDEVIATION | V12 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5122 | 35123 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V23_OVERDEVIATION | V23 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5124 | 35125 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V31_OVERDEVIATION | V31 average over deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5126 | 35127 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_VN_UNDERDEVIATION | VN average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5128 | 35129 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V1_UNDERDEVIATION | V1 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5130 | 35131 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V2_UNDERDEVIATION | V2 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5132 | 35133 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V3_UNDERDEVIATION | V3 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5134 | 35135 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V12_UNDERDEVIATION | V12 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5136 | 35137 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V23_UNDERDEVIATION | V23 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5138 | 35139 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_3SEC_V31_UNDERDEVIATION | V31 average under deviation calculated over 3 seconds by 10/12 cycles values | 4 | R | 4 Bytes Float | 5140 | 35141 | Power Quality Analyzer G4400
PRM_CODE_AVG_FLAG_10MINUTES | Indicate an trigger for event, within the last checked 10 minutes time interval | 4 | R | 4 Bytes Integer | 5142 | 35143 | Power Quality Analyzer G4400
PRM_CODE_AVG_TIMESTAMP_10MINUTES | Start Time of the 10 minutes interval (at a 100 nsec steps accuracy) | 8 | R | 8 Bytes FileTime | 5144 | 35145 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_VN_RMS | VN RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5148 | 35149 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V1_RMS | V1 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5150 | 35151 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V2_RMS | V2 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5152 | 35153 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V3_RMS | V3 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5154 | 35155 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V12_RMS | V12 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5156 | 35157 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V23_RMS | V23 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5158 | 35159 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V31_RMS | V31 RMS aggregated value over 10 minutes, aggregation done on 10/12 cycles intervals. | 4 | R | 4 Bytes Float | 5160 | 35161 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V_UNBALANCE_ZERO_RATIO | 10 minutes averaging - u0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5162 | 35163 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_I_UNBALANCE_ZERO_RATIO | 10 minutes averaging - i0 Zero sequence ratio | 4 | R | 4 Bytes Float | 5164 | 35165 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V_UNBALANCE | V average unbalance calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5166 | 35167 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_I_UNBALANCE | I average unbalance calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5168 | 35169 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V_POSITIVE_UNBALANCE | V average sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5170 | 35171 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_I_POSITIVE_UNBALANCE | I average positive sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5172 | 35173 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V_NEGATIVE_UNBALANCE | V average negative sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5174 | 35175 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_I_NEGATIVE_UNBALANCE | I average negative sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5176 | 35177 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V_ZERO_UNBALANCE | V average zero sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5178 | 35179 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_I_ZERO_UNBALANCE | I average zero sequence calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5180 | 35181 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_VN_OVERDEVIATION | VN average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5182 | 35183 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V1_OVERDEVIATION | V1 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5184 | 35185 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V2_OVERDEVIATION | V2 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5186 | 35187 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V3_OVERDEVIATION | V3 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5188 | 35189 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V12_OVERDEVIATION | V12 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5190 | 35191 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V23_OVERDEVIATION | V23 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5192 | 35193 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V31_OVERDEVIATION | V31 average over deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5194 | 35195 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_VN_UNDERDEVIATION | VN average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5196 | 35197 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V1_UNDERDEVIATION | V1 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5198 | 35199 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V2_UNDERDEVIATION | V2 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5200 | 35201 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V3_UNDERDEVIATION | V3 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5202 | 35203 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V12_UNDERDEVIATION | V12 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5204 | 35205 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V23_UNDERDEVIATION | V23 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5206 | 35207 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_10MIN_V31_UNDERDEVIATION | V31 average under deviation calculated over 10 minutes by 10/12 cycles values | 4 | R | 4 Bytes Float | 5208 | 35209 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_THD | VN THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3401 | 33402 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_THD | V1 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3403 | 33404 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_THD | V2 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3405 | 33406 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_THD | V3 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3407 | 33408 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_THD | V12 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3409 | 33410 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_THD | V23 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3411 | 33412 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_THD | V31 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3413 | 33414 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_THD | IN THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3415 | 33416 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_THD | IDC THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3417 | 33418 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_THD | I1 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3419 | 33420 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_THD | I2 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3421 | 33422 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_THD | I3 THD, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3423 | 33424 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_THD_EVEN | VN THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2544 | 32545 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_THD_EVEN | V1 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2546 | 32547 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_THD_EVEN | V2 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2548 | 32549 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_THD_EVEN | V3 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2550 | 32551 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_THD_EVEN | V12 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2552 | 32553 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_THD_EVEN | V23 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2554 | 32555 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_THD_EVEN | V31 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2556 | 32557 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_THD_EVEN | IN THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2558 | 32559 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_THD_EVEN | IDC THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2560 | 32561 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_THD_EVEN | I1 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2562 | 32563 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_THD_EVEN | I2 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2564 | 32565 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_THD_EVEN | I3 THD even, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2566 | 32567 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_THD_ODD | VN THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2568 | 32569 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_THD_ODD | V1 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2570 | 32571 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_THD_ODD | V2 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2572 | 32573 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_THD_ODD | V3 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2574 | 32575 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_THD_ODD | V12 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2576 | 32577 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_THD_ODD | V23 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2578 | 32579 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_THD_ODD | V31 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2580 | 32581 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_THD_ODD | IN THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2582 | 32583 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_THD_ODD | IDC THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2584 | 32585 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_THD_ODD | I1 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2586 | 32587 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_THD_ODD | I2 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2588 | 32589 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_THD_ODD | I3 THD odd, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 2590 | 32591 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_N | LN active power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3425 | 33426 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_1 | Line 1 active power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3427 | 33428 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_2 | Line 2 active power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3429 | 33430 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_3 | Line 3 active power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3431 | 33432 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2_3_THD_AVERAGE |  | 4 | R | 4 Bytes Float | 3433 | 33434 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_23_31_THD_AVERAGE |  | 4 | R | 4 Bytes Float | 3435 | 33436 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_SUM | Total Active Power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3439 | 33440 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_N | LN Reactive power (VAr), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3441 | 33442 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_1 | Line 1 Reactive power (VAr) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3443 | 33444 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_2 | Line 2 Reactive power (VAr) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3445 | 33446 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_3 | Line 3 Reactive power (VAr) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3447 | 33448 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_SUM | Reactive power - Total (VAr) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3455 | 33456 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_N | LN apparent power (VA), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3457 | 33458 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_1 | Line 1 apparent power (VA) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3459 | 33460 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_2 | Line 2 apparent power (VA) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3461 | 33462 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_3 | Line 3 apparent power (VA) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3463 | 33464 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_SUM | Apparent Power Total (VA) calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3471 | 33472 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_N | Power factor N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3473 | 33474 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_1 | Power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3475 | 33476 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_2 | Power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3477 | 33478 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_3 | Power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3479 | 33480 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_TOTAL | Total Power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3481 | 33482 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_1 | Fundamental Power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3483 | 33484 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_2 | Power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3485 | 33486 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_3 | Fundamental Power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3487 | 33488 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_TOTAL | Fundamental Total Power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3489 | 33490 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_N_RANGE |  | 4 | R | 4 Bytes Float | 4167 | 34168 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_1_RANGE |  | 4 | R | 4 Bytes Float | 4169 | 34170 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_2_RANGE |  | 4 | R | 4 Bytes Float | 4171 | 34172 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_3_RANGE |  | 4 | R | 4 Bytes Float | 4173 | 34174 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_TOTAL_RANGE |  | 4 | R | 4 Bytes Float | 4175 | 34176 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_1_RANGE |  | 4 | R | 4 Bytes Float | 4177 | 34178 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_2_RANGE |  | 4 | R | 4 Bytes Float | 4179 | 34180 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_3_RANGE |  | 4 | R | 4 Bytes Float | 4181 | 34182 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_TOTAL_RANGE |  | 4 | R | 4 Bytes Float | 4183 | 34184 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_RMS | VN RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3503 | 33504 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_RMS | V1 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3505 | 33506 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_RMS | V2 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3507 | 33508 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_RMS | V3 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3509 | 33510 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_RMS | V12 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3511 | 33512 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_RMS | V23 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3513 | 33514 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_RMS | V31 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3515 | 33516 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_RMS | IN RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3517 | 33518 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_RMS | IDC RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3519 | 33520 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_RMS | I1 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3521 | 33522 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_RMS | I2 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3523 | 33524 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_RMS | I3 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3525 | 33526 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2_3_RMS_AVERAGE | Average RMS Line Voltage , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3527 | 33528 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_23_31_RMS_AVERAGE | Average RMS phase to phase Voltage , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3529 | 33530 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_2_3_RMS_AVERAGE | Average RMS Current , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3531 | 33532 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_CREST_FACTOR | VN Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3533 | 33534 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_CREST_FACTOR | V1 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3535 | 33536 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_CREST_FACTOR | V2 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3537 | 33538 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_CREST_FACTOR | V3 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3539 | 33540 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_CREST_FACTOR | V12 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3541 | 33542 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_CREST_FACTOR | V23 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3543 | 33544 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_CREST_FACTOR | V31 Crest factor, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3545 | 33546 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_CREST_FACTOR | IN Crest factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3547 | 33548 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_CREST_FACTOR | IDC Crest factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3549 | 33550 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_CREST_FACTOR | I1 Crest factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3551 | 33552 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_CREST_FACTOR | I2 Crest factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3553 | 33554 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_CREST_FACTOR | I3 Crest factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3555 | 33556 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_K_FACTOR | IN K factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3557 | 33558 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_K_FACTOR | IDC K factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3559 | 33560 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_K_FACTOR | I1 K factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3561 | 33562 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_K_FACTOR | I2 K factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3563 | 33564 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_K_FACTOR | I3 K factor, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3565 | 33566 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_POSITIVE_UNBALANCE | Positive sequence Voltage, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2592 | 32593 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_POSITIVE_UNBALANCE | Positive sequence Current, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2594 | 32595 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_NEGATIVE_UNBALANCE | Negative sequence voltage, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2596 | 32597 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_NEGATIVE_UNBALANCE | Negative sequence current, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2598 | 32599 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_ZERO_UNBALANCE | Zero sequence Voltage, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2600 | 32601 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_ZERO_UNBALANCE | Zero sequence Current, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 2602 | 32603 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_UNBALANCE_ZERO_RATIO | [AVERAGE] u0 Zero sequence ratio | 4 | R | 4 Bytes Float | 2604 | 32605 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_UNBALANCE_ZERO_RATIO | [AVERAGE] i0 Zero sequence ratio | 4 | R | 4 Bytes Float | 2606 | 32607 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_UNBALANCE | Voltage unbalance, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2608 | 32609 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_UNBALANCE | Current unbalance, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2610 | 32611 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_OVERDEVIATION | VN over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2612 | 32613 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_OVERDEVIATION | V1 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2614 | 32615 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_OVERDEVIATION | V2 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2616 | 32617 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_OVERDEVIATION | V3 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2618 | 32619 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_OVERDEVIATION | V12 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2620 | 32621 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_OVERDEVIATION | V23 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2622 | 32623 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_OVERDEVIATION | V31 over deviation calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2624 | 32625 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_UNDERDEVIATION | VN under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2626 | 32627 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_UNDERDEVIATION | V1 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2628 | 32629 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_UNDERDEVIATION | V2 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2630 | 32631 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_UNDERDEVIATION | V3 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2632 | 32633 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_UNDERDEVIATION | V12 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2634 | 32635 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_UNDERDEVIATION | V23 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2636 | 32637 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_UNDERDEVIATION | V31 under deviation, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 2638 | 32639 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_VN_USR | Average VN over 10/12 cycles | 4 | R | 4 Bytes Float | 3871 | 33872 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V1_USR | Average V1 over 10/12 cycles | 4 | R | 4 Bytes Float | 3873 | 33874 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V2_USR | Average V2 over 10/12 cycles | 4 | R | 4 Bytes Float | 3875 | 33876 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V3_USR | Average V3 over 10/12 cycles | 4 | R | 4 Bytes Float | 3877 | 33878 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V12_USR | Average V12 over 10/12 cycles | 4 | R | 4 Bytes Float | 3879 | 33880 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V23_USR | Average V23 over 10/12 cycles | 4 | R | 4 Bytes Float | 3881 | 33882 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V31_USR | Average V31 over 10/12 cycles | 4 | R | 4 Bytes Float | 3883 | 33884 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_VN_THD | Average THD VN over 10/12 cycles | 4 | R | 4 Bytes Float | 3893 | 33894 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V1_THD | Average THD V1 over 10/12 cycles | 4 | R | 4 Bytes Float | 3895 | 33896 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V2_THD | Average THD V2 over 10/12 cycles | 4 | R | 4 Bytes Float | 3897 | 33898 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V3_THD | Average THD V3 over 10/12 cycles | 4 | R | 4 Bytes Float | 3899 | 33900 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V12_THD | Average THD V12 over 10/12 cycles | 4 | R | 4 Bytes Float | 3901 | 33902 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V23_THD | Average THD V23 over 10/12 cycles | 4 | R | 4 Bytes Float | 3903 | 33904 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V31_THD | Average THD V31 over 10/12 cycles | 4 | R | 4 Bytes Float | 3905 | 33906 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_IN_THD | Average THD IN over 10/12 cycles | 4 | R | 4 Bytes Float | 3907 | 33908 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I1_THD | Average THD I1 over 10/12 cycles | 4 | R | 4 Bytes Float | 3909 | 33910 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I2_THD | Average THD I2 over 10/12 cycles | 4 | R | 4 Bytes Float | 3911 | 33912 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I3_THD | Average THD I3 over 10/12 cycles | 4 | R | 4 Bytes Float | 3913 | 33914 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_ACTIVE_POWER_N | Average active power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3915 | 33916 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_ACTIVE_POWER_MAIN_1 | Average active power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3917 | 33918 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_ACTIVE_POWER_MAIN_2 | Average active power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3919 | 33920 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_ACTIVE_POWER_MAIN_3 | Average active power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3921 | 33922 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_ACTIVE_POWER_SUM | Average total active power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3929 | 33930 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_REACTIVE_POWER_N | Average reactive power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3931 | 33932 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_REACTIVE_POWER_MAIN_1 | Average reactive power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3933 | 33934 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_REACTIVE_POWER_MAIN_2 | Average reactive power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3935 | 33936 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_REACTIVE_POWER_MAIN_3 | Average reactive power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3937 | 33938 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_REACTIVE_POWER_SUM | Average total reactive power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3945 | 33946 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_APPARENT_POWER_N | Average apparent power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3947 | 33948 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_APPARENT_POWER_MAIN_1 | Average apparent power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3949 | 33950 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_APPARENT_POWER_MAIN_2 | Average apparent power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3951 | 33952 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_APPARENT_POWER_MAIN_3 | Average apparent power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3953 | 33954 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_APPARENT_POWER_SUM | Average total apparent power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3961 | 33962 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_POWER_FACTOR_N | Average power factor N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3963 | 33964 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_POWER_FACTOR_1 | Average power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3965 | 33966 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_POWER_FACTOR_2 | Average power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3967 | 33968 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_POWER_FACTOR_3 | Average power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3969 | 33970 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_POWER_FACTOR_TOTAL | Average total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3971 | 33972 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_PF_FUNDAMENTAL_1 | Average fundamental power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3973 | 33974 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_PF_FUNDAMENTAL_2 | Average fundamental power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3975 | 33976 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_PF_FUNDAMENTAL_3 | Average fundamental power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3977 | 33978 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_PF_FUNDAMENTAL_TOTAL | Average fundamental total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3979 | 33980 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_VN_RMS | VN RMS average reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3981 | 33982 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V1_RMS | V1 RMS average reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3983 | 33984 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V2_RMS | V2 average RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3985 | 33986 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V3_RMS | V3 average RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3987 | 33988 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V12_RMS | V12 average RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3989 | 33990 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V23_RMS | V23 average RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3991 | 33992 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V31_RMS | V31 average RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3993 | 33994 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_IN_RMS | IN average RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3995 | 33996 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I1_RMS | I1 average RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3997 | 33998 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I2_RMS | I2 average RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3999 | 34000 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I3_RMS | I3 average RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 4001 | 34002 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V1_2_3_RMS_AVERAGE | Average RMS Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 4003 | 34004 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V12_23_31_RMS_AVERAGE | Average RMS phase to phase Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 4005 | 34006 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I1_2_3_RMS_AVERAGE | Average RMS Current, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 4007 | 34008 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_VN_CREST_FACTOR | Average VN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4009 | 34010 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V1_CREST_FACTOR | Average V1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4011 | 34012 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V2_CREST_FACTOR | Average V2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4013 | 34014 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V3_CREST_FACTOR | Average V3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4015 | 34016 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V12_CREST_FACTOR | Average V12 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4017 | 34018 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V23_CREST_FACTOR | Average V23 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4019 | 34020 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V31_CREST_FACTOR | Average V31 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4021 | 34022 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_IN_CREST_FACTOR | Average IN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4023 | 34024 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I1_CREST_FACTOR | Average I1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4025 | 34026 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I2_CREST_FACTOR | Average I2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4027 | 34028 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I3_CREST_FACTOR | Average I3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4029 | 34030 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_IN_K_FACTOR | Average IN K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4031 | 34032 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I1_K_FACTOR | Average I1 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4033 | 34034 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I2_K_FACTOR | Average I2 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4035 | 34036 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I3_K_FACTOR | Average I3 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4037 | 34038 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_V_UNBALANCE | Voltage average unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4039 | 34040 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_I_UNBALANCE | Current average unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 4041 | 34042 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_THD_MIN | Minimum VN THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3567 | 33568 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_THD_MIN | Minimum V1 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3569 | 33570 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_THD_MIN | Minimum V2 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3571 | 33572 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_THD_MIN | Minimum V3 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3573 | 33574 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_THD_MIN | Minimum V12 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3575 | 33576 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_THD_MIN | Minimum V23 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3577 | 33578 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_THD_MIN | Minimum V31 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3579 | 33580 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_THD_MIN | Minimum IN THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3581 | 33582 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_THD_MIN | Minimum I1 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3583 | 33584 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_THD_MIN | Minimum I2 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3585 | 33586 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_THD_MIN | Minimum I3 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3587 | 33588 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_N_MIN | Minimum N Active power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3589 | 33590 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_1_MIN | Minimum active power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3591 | 33592 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_2_MIN | Minimum active power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3593 | 33594 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_3_MIN | Minimum active power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3595 | 33596 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_SUM_MIN | Minimum total active power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3603 | 33604 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_N_MIN | Minimum reactive power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3605 | 33606 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_1_MIN | Minimum reactive power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3607 | 33608 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_2_MIN | Minimum reactive power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3609 | 33610 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_3_MIN | Minimum reactive power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3611 | 33612 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_SUM_MIN | Minimum total reactive power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3619 | 33620 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_N_MIN | Minimum apparent power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3621 | 33622 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_1_MIN | Minimum apparent power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3623 | 33624 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_2_MIN | Minimum apparent power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3625 | 33626 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_3_MIN | Minimum apparent power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3627 | 33628 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_SUM_MIN | Minimum total apparent power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3635 | 33636 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_N_MIN | Minimum power factor N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3637 | 33638 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_1_MIN | Minimum power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3639 | 33640 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_2_MIN | Minimum power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3641 | 33642 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_3_MIN | Minimum power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3643 | 33644 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_TOTAL_MIN | Minimum total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3645 | 33646 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_1_MIN | Minimum fundamental power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3647 | 33648 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_2_MIN | Minimum fundamental power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3649 | 33650 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_3_MIN | Minimum fundamental power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3651 | 33652 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_TOTAL_MIN | Minimum fundamental total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3653 | 33654 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_RMS_MIN | VN RMS Minimum reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3655 | 33656 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_RMS_MIN | V1 RMS Minimum reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3657 | 33658 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_RMS_MIN | V2 Minimum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3659 | 33660 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_RMS_MIN | V3 Minimum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3661 | 33662 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_RMS_MIN | V12 Minimum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3663 | 33664 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_RMS_MIN | V23 Minimum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3665 | 33666 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_RMS_MIN | V31 Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3667 | 33668 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_RMS_MIN | IN Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3669 | 33670 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_RMS_MIN | IDC Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3671 | 33672 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_RMS_MIN | I1 Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3673 | 33674 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_RMS_MIN | I2 Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3675 | 33676 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_RMS_MIN | I3 Minimum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3677 | 33678 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2_3_RMS_AVERAGE_MIN | Minimum RMS Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3679 | 33680 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_23_31_RMS_AVERAGE_MIN | Minimum RMS phase to phase Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3681 | 33682 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_2_3_RMS_AVERAGE_MIN | Minimum RMS Current, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3683 | 33684 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_CREST_FACTOR_MIN | Minimum VN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3685 | 33686 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_CREST_FACTOR_MIN | Minimum V1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3687 | 33688 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_CREST_FACTOR_MIN | Minimum V2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3689 | 33690 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_CREST_FACTOR_MIN | Minimum V3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3691 | 33692 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_CREST_FACTOR_MIN | Minimum V12 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3693 | 33694 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_CREST_FACTOR_MIN | Minimum V23 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3695 | 33696 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_CREST_FACTOR_MIN | Minimum V31 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3697 | 33698 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_CREST_FACTOR_MIN | Minimum IN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3699 | 33700 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_CREST_FACTOR_MIN | Minimum I1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3701 | 33702 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_CREST_FACTOR_MIN | Minimum I2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3703 | 33704 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_CREST_FACTOR_MIN | Minimum I3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3705 | 33706 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_K_FACTOR_MIN | Minimum IN K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3707 | 33708 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_K_FACTOR_MIN | Minimum I1 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3709 | 33710 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_K_FACTOR_MIN | Minimum I2 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3711 | 33712 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_K_FACTOR_MIN | Minimum I3 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3713 | 33714 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_UNBALANCE_MIN | Voltage Minimum unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3715 | 33716 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_UNBALANCE_MIN | Current Minimum unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3717 | 33718 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_THD_MAX | Maximum VN THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3719 | 33720 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_THD_MAX | Maximum V1 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3721 | 33722 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_THD_MAX | Maximum V2 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3723 | 33724 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_THD_MAX | Maximum V3 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3725 | 33726 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_THD_MAX | Maximum V12 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3727 | 33728 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_THD_MAX | Maximum V23 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3729 | 33730 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_THD_MAX | Maximum V31 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3731 | 33732 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_THD_MAX | Maximum IN THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3733 | 33734 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_THD_MAX | Maximum I1 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3735 | 33736 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_THD_MAX | Maximum I2 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3737 | 33738 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_THD_MAX | Maximum I3 THD, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3739 | 33740 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_N_MAX | Maximum N Active power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3741 | 33742 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_1_MAX | Maximum active power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3743 | 33744 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_2_MAX | Maximum active power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3745 | 33746 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_MAIN_3_MAX | Maximum active power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3747 | 33748 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_SUM_MAX | Maximum total active power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3755 | 33756 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_N_MAX | Maximum reactive power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3757 | 33758 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_1_MAX | Maximum reactive power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3759 | 33760 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_2_MAX | Maximum reactive power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3761 | 33762 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_MAIN_3_MAX | Maximum reactive power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3763 | 33764 | Power Quality Analyzer G4400
PRM_CODE_LONG_REACTIVE_POWER_SUM_MAX | Maximum total reactive power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3771 | 33772 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_N_MAX | Maximum apparent power N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3773 | 33774 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_1_MAX | Maximum apparent power phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3775 | 33776 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_2_MAX | Maximum apparent power phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3777 | 33778 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_MAIN_3_MAX | Maximum apparent power phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3779 | 33780 | Power Quality Analyzer G4400
PRM_CODE_LONG_APPARENT_POWER_SUM_MAX | Maximum total apparent power, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3787 | 33788 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_N_MAX | Maximum power factor N, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3789 | 33790 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_1_MAX | Maximum power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3791 | 33792 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_2_MAX | Maximum power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3793 | 33794 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_3_MAX | Maximum power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3795 | 33796 | Power Quality Analyzer G4400
PRM_CODE_LONG_POWER_FACTOR_TOTAL_MAX | Maximum total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3797 | 33798 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_1_MAX | Maximum fundamental power factor phase 1, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3799 | 33800 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_2_MAX | Maximum fundamental power factor phase 2, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3801 | 33802 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_3_MAX | Maximum fundamental power factor phase 3, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3803 | 33804 | Power Quality Analyzer G4400
PRM_CODE_LONG_PF_FUNDAMENTAL_TOTAL_MAX | Maximum fundamental total power factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3805 | 33806 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_RMS_MAX | VN RMS Maximum reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3807 | 33808 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_RMS_MAX | V1 RMS Maximum reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3809 | 33810 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_RMS_MAX | V2 Maximum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3811 | 33812 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_RMS_MAX | V3 Maximum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3813 | 33814 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_RMS_MAX | V12 Maximum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3815 | 33816 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_RMS_MAX | V23 Maximum RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3817 | 33818 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_RMS_MAX | V31 Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3819 | 33820 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_RMS_MAX | IN Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3821 | 33822 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_RMS_MAX | IDC Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3823 | 33824 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_RMS_MAX | I1 Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3825 | 33826 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_RMS_MAX | I2 Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3827 | 33828 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_RMS_MAX | I3 Maximum RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3829 | 33830 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2_3_RMS_AVERAGE_MAX | Maximum RMS Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3831 | 33832 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_23_31_RMS_AVERAGE_MAX | Maximum RMS phase to phase Voltage, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3833 | 33834 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_2_3_RMS_AVERAGE_MAX | Maximum RMS Current, calculated every 10/12 cycles | 4 | R | 4 Bytes Float | 3835 | 33836 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_CREST_FACTOR_MAX | Minimum VN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3837 | 33838 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_CREST_FACTOR_MAX | Minimum V1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3839 | 33840 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_CREST_FACTOR_MAX | Minimum V2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3841 | 33842 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_CREST_FACTOR_MAX | Minimum V3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3843 | 33844 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_CREST_FACTOR_MAX | Minimum V12 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3845 | 33846 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_CREST_FACTOR_MAX | Minimum V23 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3847 | 33848 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_CREST_FACTOR_MAX | Minimum V31 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3849 | 33850 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_CREST_FACTOR_MAX | Minimum IN crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3851 | 33852 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_CREST_FACTOR_MAX | Minimum I1 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3853 | 33854 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_CREST_FACTOR_MAX | Minimum I2 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3855 | 33856 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_CREST_FACTOR_MAX | Minimum I3 crest factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3857 | 33858 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_K_FACTOR_MAX | Minimum IN K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3859 | 33860 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_K_FACTOR_MAX | Minimum I1 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3861 | 33862 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_K_FACTOR_MAX | Minimum I2 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3863 | 33864 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_K_FACTOR_MAX | Minimum I3 K factor, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3865 | 33866 | Power Quality Analyzer G4400
PRM_CODE_LONG_V_UNBALANCE_MAX | Voltage Minimum unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3867 | 33868 | Power Quality Analyzer G4400
PRM_CODE_LONG_I_UNBALANCE_MAX | Current Minimum unbalance, calculated over 10/12 cycles | 4 | R | 4 Bytes Float | 3869 | 33870 | Power Quality Analyzer G4400
PRM_CODE_LOG_REPORT_FILTER |  | 2 | RW | 2 Bytes Integer | 225 | 30226 | Power Quality Analyzer G4400
PRM_CODE_LOG_REPORT_CREATE_FILE | Export log to csv file (the file is created on the CF) [Not an electrical parameter!!!!!] | 0 | RW | Action | 226 | 30227 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IN_FLP | Total Active energy in since last reset (W) in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2167 | 32168 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VA_IN_FLP | Total Apparent energy In since last reset (VA) in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2169 | 32170 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IN_FLP | Total Reactive energy out since last reset (VAr) in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2171 | 32172 | Power Quality Analyzer G4400
PRM_CODE_METERING_W_IN_FLP | … in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2173 | 32174 | Power Quality Analyzer G4400
PRM_CODE_METERING_VA_IN_FLP | … in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2175 | 32176 | Power Quality Analyzer G4400
PRM_CODE_METERING_VAr_IN_FLP | … in floating point format (4Bytes) | 4 | R | 4 Bytes Float | 2177 | 32178 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IPO_FLP | Total energy (W) in + Total energy (W) out | 4 | R | 4 Bytes Float | 2179 | 32180 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IPO_FLP | Total energy (WAr) in + Total energy (WAr) out | 4 | R | 4 Bytes Float | 2181 | 32182 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_W_IMO_FLP | Total energy (W) in Minus Total energy (W) out | 4 | R | 4 Bytes Float | 2183 | 32184 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_VAr_IMO_FLP | Total energy (WAr) in Minus Total energy (WAr) out | 4 | R | 4 Bytes Float | 2185 | 32186 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IPO_FLP |  | 4 | R | 4 Bytes Float | 2193 | 32194 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IPO_FLP |  | 4 | R | 4 Bytes Float | 2195 | 32196 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IMO_FLP |  | 4 | R | 4 Bytes Float | 2197 | 32198 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IMO_FLP |  | 4 | R | 4 Bytes Float | 2199 | 32200 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_IN_FLP |  | 4 | R | 4 Bytes Float | 2201 | 32202 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_W_OUT_FLP |  | 4 | R | 4 Bytes Float | 2203 | 32204 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VA_IN_FLP |  | 4 | R | 4 Bytes Float | 2205 | 32206 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_IN_FLP |  | 4 | R | 4 Bytes Float | 2207 | 32208 | Power Quality Analyzer G4400
PRM_CODE_TOTAL_ENERGY_2_VAr_OUT_FLP |  | 4 | R | 4 Bytes Float | 2209 | 32210 | Power Quality Analyzer G4400

## Лист1
PRM_CODE_LONG_VN_RMS | VN RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3503 | 33504 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_RMS | V1 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3505 | 33506 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_RMS | V2 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3507 | 33508 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_RMS | V3 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3509 | 33510 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_RMS | V12 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3511 | 33512 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_RMS | V23 RMS reading, calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3513 | 33514 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_RMS | V31 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3515 | 33516 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_RMS | I1 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3521 | 33522 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_RMS | I2 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3523 | 33524 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_RMS | I3 RMS reading, calculated every 10/12 cycles (A) | 4 | R | 4 Bytes Float | 3525 | 33526 | Power Quality Analyzer G4400
PRM_CODE_LONG_ACTIVE_POWER_SUM | Total Active Power (W), calculate every 10/12 cycles | 4 | R | 4 Bytes Float | 3439 | 33440 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2_3_RMS_AVERAGE | Average RMS Line Voltage , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3527 | 33528 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_23_31_RMS_AVERAGE | Average RMS phase to phase Voltage , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3529 | 33530 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_2_3_RMS_AVERAGE | Average RMS Current , calculated every 10/12 cycles (v) | 4 | R | 4 Bytes Float | 3531 | 33532 | Power Quality Analyzer G4400

## Device Logs
Parameter | Description | Byte Size | R/W | Type | MODBUS Code | MODBUS Code With Offset | Supported Device
PRM_CODE_METERING_ENERGY_INTERVAL | User configurated time interval for the energy meter | 4 | RW | 4 Bytes Integer | 352 | 30353 | Power Quality Analyzer G4400
PRM_CODE_METER_READ_LOG_STATE | Set the active reports (0 - disable; 1 - energy; 2 - parameters; 3- Set both of them) | 4 | RW | 4 Bytes Integer | 342 | 30343 | Power Quality Analyzer G4401
PRM_CODE_METER_READ_FILE_RATE | Set the rate of reports: 0 - day; 1- week; 2- month | 4 | RW | 4 Bytes Integer | 344 | 30345 | Power Quality Analyzer G4402
PRM_CODE_LOG_REPORT_TIMESTAMP_START | Start of the requested period to get logs (at 100nsec steps accuracy) | 8 | RW | 8 Bytes FileTime | 332 | 30333 | Power Quality Analyzer G4403
PRM_CODE_LOG_REPORT_TIMESTAMP_END | End of the requested period to get logs (at 100nsec steps accuracy) | 8 | RW | 8 Bytes FileTime | 336 | 30337 | Power Quality Analyzer G4404

## Configuration Parameters
Parameter | Description | Byte Size | R/W | Type | MODBUS Code | MODBUS Code With Offset | Supported Device
PRM_CODE_METER_READ_FILE_RESTART_TIME | Restart time of the report file | 4 | RW | 4 Bytes Integer | 346 | 30347 | Power Quality Analyzer G4400
PRM_CODE_METER_READ_FILE_RESTART_MINUTE | Restart minute of the report file | 4 | RW | 4 Bytes Integer | 348 | 30349 | Power Quality Analyzer G4400
PRM_CODE_METER_READ_FILE_RESTART_DAY | Restart day of the report file | 4 | RW | 4 Bytes Integer | 350 | 30351 | Power Quality Analyzer G4400
PRM_CODE_AUTHENTICATION | Authentication field for get write access through modbus | 2 | W | 2 Bytes Integer | 998 | 30999 | Power Quality Analyzer G4400
PRM_CODE_SYSTEM_FILETIME | Get system time as FILETIME | 8 | R | 8 Bytes FileTime | 375 | 30376 | Power Quality Analyzer G4400
PRM_CODE_RTC_COUNTER | Get the time passed from manufacture (in FILETIME format) | 8 | R | 8 Bytes FileTime | 369 | 30370 | Power Quality Analyzer G4400
PRM_CODE_CYCLE_TIMESTAMP_WRD |  | 8 | R | 8 Bytes FileTime | 379 | 30380 | Power Quality Analyzer G4400
PRM_CODE_SERIAL_NUMBER | Device serial number | 30 | R | String | 629 | 30630 | Power Quality Analyzer G4400
PRM_CODE_REGIONAL_TIME_ZONE | UTC time zone | 4 | RW | 4 Bytes Integer | 373 | 30374 | Power Quality Analyzer G4400
PRM_CODE_ENERGY_REPORT_CREATE_FILE | Manual create energy report using PRM_CODE_LOG_REPORT_TIMESTAMP_START and PRM_CODE_LOG_REPORT_TIMESTAMP_END as time range | 0 | RW | Action | 228 | 30229 | Power Quality Analyzer G4400
PRM_CODE_REPORTS_STATUS | Status of the manual creating energy report | 4 | RW | 4 Bytes Integer | 644 | 30645 | Power Quality Analyzer G4400
PRM_CODE_REPORTS_ERASE_ALL | Delete the contents of directory: CF/Reports/Requests | 0 | RW | Action | 330 | 30331 | Power Quality Analyzer G4400
PRM_CODE_REPORTS_STATUS_CLEAR | Clear the status (PRM_CODE_REPORTS_STATUS) bitmask | 0 | RW | Action | 340 | 30341 | Power Quality Analyzer G4400
PRM_CODE_DST_STATUS | Status of Daylight saving time | 4 | RW | 4 Bytes Integer | 359 | 30360 | Power Quality Analyzer G4400
PRM_CODE_DST_START_TIME | Start time of Daylight saving time (in FILETIME format) | 8 | RW | 8 Bytes FileTime | 361 | 30362 | Power Quality Analyzer G4400
PRM_CODE_DST_END_TIME | End time of Daylight saving time (in FILETIME format) | 8 | RW | 8 Bytes FileTime | 365 | 30366 | Power Quality Analyzer G4400
PRM_CODE_IDLE_TIMEOUT | Time out to abort command | 4 | RW | 4 Bytes Integer | 357 | 30358 | Power Quality Analyzer G4400
PRM_CODE_MAC_ADDRESS | Motherboard MAC address | 6 | R | String | 354 | 30355 | Power Quality Analyzer G4400
PRM_CODE_TIME_SYNC_STATUS | Time synchronization (excellent/good/poor…) | 4 | R | 4 Bytes Integer | 449 | 30450 | Power Quality Analyzer G4400
PRM_CODE_CF_FREE_SPACE | Compact flash space availability (Bytes) | 8 | R | 4 Bytes Integer | 451 | 30452 | Power Quality Analyzer G4400
PRM_CODE_CF_TOTAL_SPACE | Compact flash storage size | 8 | R | 4 Bytes Integer | 455 | 30456 | Power Quality Analyzer G4400
PRM_CODE_PHY1_LINK | Lan1 status (on/off) | 4 | R | 4 Bytes Integer | 493 | 30494 | Power Quality Analyzer G4400
PRM_CODE_PHY2_LINK | Lan2 status (on/off) | 4 | R | 4 Bytes Integer | 495 | 30496 | Power Quality Analyzer G4400
PRM_CODE_PHY1_SPEED | LAN1 communication speed | 4 | R | 4 Bytes Integer | 497 | 30498 | Power Quality Analyzer G4400
PRM_CODE_PHY2_SPEED | LAN2 communication speed | 4 | R | 4 Bytes Integer | 499 | 30500 | Power Quality Analyzer G4400
PRM_CODE_PHY1_MODE | LAN1 negotiation method | 4 | RW | 4 Bytes Integer | 37 | 30038 | Power Quality Analyzer G4400
PRM_CODE_PHY2_MODE | LAN2 negotiation method (100half,100full, 10half,10full, auto) | 4 | RW | 4 Bytes Integer | 39 | 30040 | Power Quality Analyzer G4400
PRM_CODE_SW_VERSION |  |  | R | String | 419 | 30420 | Power Quality Analyzer G4400
PRM_CODE_PHY1_DUPLEX | Half / full duplex | 4 | R | 4 Bytes Integer | 501 | 30502 | Power Quality Analyzer G4400
PRM_CODE_PHY2_DUPLEX | Half / full duplex | 4 | R | 4 Bytes Integer | 503 | 30504 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_CONNECTIONS_WEB | number of current active web interfaces | 4 | R | 4 Bytes Integer | 505 | 30506 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_CONNECTIONS_OPC | Number of current OPC clients | 4 | R | 4 Bytes Integer | 507 | 30508 | Power Quality Analyzer G4400
PRM_CODE_ACTIVE_CONNECTIONS_LCD | Number of LCD's active | 4 | R | 4 Bytes Integer | 509 | 30510 | Power Quality Analyzer G4400
PRM_CODE_FTP_ACTIVE_CONNECTION | Number of active FTP clients | 4 | R | 4 Bytes Integer | 511 | 30512 | Power Quality Analyzer G4400
PRM_CODE_FTP_MAX_CONNECTION | Maximum number of FTP users at the same time | 4 | R | 4 Bytes Integer | 515 | 30516 | Power Quality Analyzer G4400
PRM_CODE_SNTP_SOURCE | The current time source (SNTP,GPS…) | 4 | R | 4 Bytes Integer | 517 | 30518 | Power Quality Analyzer G4400
PRM_CODE_SMTP_FROM_ADRS | E-mail address which used in "form:" field when sending E-Mail | 80 | RW | String | 41 | 30042 | Power Quality Analyzer G4400
PRM_CODE_SMTP_TO_ADRS | Destination address for E-Mail alert | 80 | RW | String | 81 | 30082 | Power Quality Analyzer G4400
PRM_CODE_SMTP_EVENTS | List of events which can cause E-Mail alert | 80 | RW | 2 Bytes Integer | 121 | 30122 | Power Quality Analyzer G4400
PRM_CODE_SMTP_REOCCURRENCE_TIME_LIMT | Minimum time between reoccurrence/sequence E-Mail | 4 | RW | 4 Bytes Integer | 519 | 30520 | Power Quality Analyzer G4400
PRM_CODE_POWER_POWERED_220 | Is the device powered with 220VAC? | 4 | R | 4 Bytes Integer | 521 | 30522 | Power Quality Analyzer G4400
PRM_CODE_POWER_POWERED_LAN | Is the device powered with POE? | 4 | R | 4 Bytes Integer | 523 | 30524 | Power Quality Analyzer G4400
PRM_CODE_POWER_POWERED_48 | Is the device powered with 48VDC? | 4 | R | 4 Bytes Integer | 525 | 30526 | Power Quality Analyzer G4400
PRM_CODE_POWER_DOWN | If TRUE then shutdown process started | 4 | R | 4 Bytes Integer | 527 | 30528 | Power Quality Analyzer G4400
PRM_CODE_POWER_POWERED_STATUS | Status of power source (48Vdc/230Vac/POE..) | 4 | R | 4 Bytes Integer | 529 | 30530 | Power Quality Analyzer G4400
PRM_CODE_POWER_PSE_STATUS | POE out status (enable/disable) | 4 | RW | 4 Bytes Integer | 531 | 30532 | Power Quality Analyzer G4400
PRM_CODE_POWER_PSE_ACTIVE | Test if POE out is active (true/false) | 4 | R | 4 Bytes Integer | 533 | 30534 | Power Quality Analyzer G4400
PRM_CODE_POWER_PSE_FAILURE | True if POE out has failed and need to be restarted | 4 | R | 4 Bytes Integer | 535 | 30536 | Power Quality Analyzer G4400
PRM_CODE_NOMINAL_V | User configured nominal Voltage | 4 | RW | 4 Bytes Float | 1 | 30002 | Power Quality Analyzer G4400
PRM_CODE_NOMINAL_V_DIFF | User configured nominal Voltage-to-voltage (as can configured on the web interface) | 4 | RW | 4 Bytes Float | 5 | 30006 | Power Quality Analyzer G4400
PRM_CODE_NOMINAL_FREQUENCY | User configured nominal frequency | 4 | RW | 4 Bytes Float | 9 | 30010 | Power Quality Analyzer G4400
PRM_CODE_POWER_CONFIGURATION | Network topology (Delta/star/single phase…) | 4 | RW | 4 Bytes Integer | 11 | 30012 | Power Quality Analyzer G4400
PRM_CODE_PT_PRIMARY | Potential Transformer Primary | 4 | RW | 4 Bytes Integer | 17 | 30018 | Power Quality Analyzer G4400
PRM_CODE_PT_SECONDARY | Potential Transformer Secondary | 4 | RW | 4 Bytes Integer | 19 | 30020 | Power Quality Analyzer G4400
PRM_CODE_VN_POLARITY | VN polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 21 | 30022 | Power Quality Analyzer G4400
PRM_CODE_V1_POLARITY | V1 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 23 | 30024 | Power Quality Analyzer G4400
PRM_CODE_V2_POLARITY | V2 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 25 | 30026 | Power Quality Analyzer G4400
PRM_CODE_V3_POLARITY | V3 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 27 | 30028 | Power Quality Analyzer G4400
PRM_CODE_IN_POLARITY | IN polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 29 | 30030 | Power Quality Analyzer G4400
PRM_CODE_I1_POLARITY | I1 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 31 | 30032 | Power Quality Analyzer G4400
PRM_CODE_I2_POLARITY | I2 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 33 | 30034 | Power Quality Analyzer G4400
PRM_CODE_I3_POLARITY | I3 polarity direction (normal/reverse) | 4 | RW | 4 Bytes Integer | 35 | 30036 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_RESTART | Restart the compliance mechanism | 0 | RW | Action | 221 | 30222 | Power Quality Analyzer G4400
PRM_CODE_PQ_COMPLIANCE_TYPE | Set the type of compliance (after that the compliance needs to be restarted) | 4 | RW | 4 Bytes Integer | 223 | 30224 | Power Quality Analyzer G4404

## Custom Prameters
Parameter | Description | Byte Size | R/W | Type | MODBUS Code | MODBUS Code With Offset | Supported Device
PRM_CODE_USER_DEFINE_0 | Custom parameters, therefore it | 4 | RW | 4 Bytes Float | 4200 | 34201 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_1 | may be defined: | 4 | RW | 4 Bytes Float | 4202 | 34203 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_2 | < Manually - by the user | 4 | RW | 4 Bytes Float | 4204 | 34205 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_3 | < Automatically - by another | 4 | RW | 4 Bytes Float | 4206 | 34207 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_4 | Scada system | 4 | RW | 4 Bytes Float | 4208 | 34209 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_5 | < Automatically - by another | 4 | RW | 4 Bytes Float | 4210 | 34211 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_6 | measurement system | 4 | RW | 4 Bytes Float | 4212 | 34213 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_7 |  | 4 | RW | 4 Bytes Float | 4214 | 34215 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_8 |  | 4 | RW | 4 Bytes Float | 4216 | 34217 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_9 |  | 4 | RW | 4 Bytes Float | 4218 | 34219 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_10 |  | 4 | RW | 4 Bytes Float | 4220 | 34221 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_11 |  | 4 | RW | 4 Bytes Float | 4222 | 34223 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_12 |  | 4 | RW | 4 Bytes Float | 4224 | 34225 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_13 |  | 4 | RW | 4 Bytes Float | 4226 | 34227 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_14 |  | 4 | RW | 4 Bytes Float | 4228 | 34229 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_15 |  | 4 | RW | 4 Bytes Float | 4230 | 34231 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_16 |  | 4 | RW | 4 Bytes Float | 4232 | 34233 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_17 |  | 4 | RW | 4 Bytes Float | 4234 | 34235 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_18 |  | 4 | RW | 4 Bytes Float | 4236 | 34237 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_19 |  | 4 | RW | 4 Bytes Float | 4238 | 34239 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_20 |  | 4 | RW | 4 Bytes Float | 4240 | 34241 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_21 |  | 4 | RW | 4 Bytes Float | 4242 | 34243 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_22 |  | 4 | RW | 4 Bytes Float | 4244 | 34245 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_23 |  | 4 | RW | 4 Bytes Float | 4246 | 34247 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_24 |  | 4 | RW | 4 Bytes Float | 4248 | 34249 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_25 |  | 4 | RW | 4 Bytes Float | 4250 | 34251 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_26 |  | 4 | RW | 4 Bytes Float | 4252 | 34253 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_27 |  | 4 | RW | 4 Bytes Float | 4254 | 34255 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_28 |  | 4 | RW | 4 Bytes Float | 4256 | 34257 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_29 |  | 4 | RW | 4 Bytes Float | 4258 | 34259 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_30 |  | 4 | RW | 4 Bytes Float | 4260 | 34261 | Power Quality Analyzer G4400
PRM_CODE_USER_DEFINE_31 |  | 4 | RW | 4 Bytes Float | 4262 | 34263 | Power Quality Analyzer G4400

## Multi IO
Parameter | Description | Byte Size | R/W | Type | MODBUS Code | MODBUS Code With Offset | Supported Device
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD1_DO1 | Relay 1 (IO module 1) | 2 | RW | 2 Bytes Integer | 4139 | 34140 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD1_DO2 | Relay 2 (IO module 1) | 2 | RW | 2 Bytes Integer | 4140 | 34141 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD1_DO3 | Relay 3 (IO module 1) | 2 | RW | 2 Bytes Integer | 4141 | 34142 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD2_DO1 | Relay 1 (IO module 2) | 2 | RW | 2 Bytes Integer | 4142 | 34143 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD2_DO2 | Relay 2 (IO module 2) | 2 | RW | 2 Bytes Integer | 4143 | 34144 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD2_DO3 | Relay 3 (IO module 2) | 2 | RW | 2 Bytes Integer | 4144 | 34145 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD3_DO1 | Relay 1 (IO module 3) | 2 | RW | 2 Bytes Integer | 4145 | 34146 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD3_DO2 | Relay 2 (IO module 3) | 2 | RW | 2 Bytes Integer | 4146 | 34147 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD3_DO3 | Relay 3 (IO module 3) | 2 | RW | 2 Bytes Integer | 4147 | 34148 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD4_DO1 | Relay 1 (IO module 4) | 2 | RW | 2 Bytes Integer | 4148 | 34149 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD4_DO2 | Relay 2 (IO module 4) | 2 | RW | 2 Bytes Integer | 4149 | 34150 | Power Quality Analyzer G4400
PRM_CODE_MIO_RELAY_WRITE_LOGIC_CARD4_DO3 | Relay 3 (IO module 4) | 2 | RW | 2 Bytes Integer | 4150 | 34151 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD1_DO1 | Digital out 1 (IO module 1) | 2 | RW | 2 Bytes Integer | 4151 | 34152 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD1_DO2 | Digital out 2 (IO module 1) | 2 | RW | 2 Bytes Integer | 4152 | 34153 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD1_DO3 | Digital out 3 (IO module 1) | 2 | RW | 2 Bytes Integer | 4153 | 34154 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD1_DO4 | Digital out 4 (IO module 1) | 2 | RW | 2 Bytes Integer | 4154 | 34155 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD2_DO1 | Digital out 1 (IO module 2) | 2 | RW | 2 Bytes Integer | 4155 | 34156 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD2_DO2 | Digital out 2 (IO module 2) | 2 | RW | 2 Bytes Integer | 4156 | 34157 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD2_DO3 | Digital out 3 (IO module 2) | 2 | RW | 2 Bytes Integer | 4157 | 34158 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD2_DO4 | Digital out 4 (IO module 2) | 2 | RW | 2 Bytes Integer | 4158 | 34159 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD3_DO1 | Digital out 1 (IO module 3) | 2 | RW | 2 Bytes Integer | 4159 | 34160 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD3_DO2 | Digital out 2 (IO module 3) | 2 | RW | 2 Bytes Integer | 4160 | 34161 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD3_DO3 | Digital out 3 (IO module 3) | 2 | RW | 2 Bytes Integer | 4161 | 34162 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD3_DO4 | Digital out 4 (IO module 3) | 2 | RW | 2 Bytes Integer | 4162 | 34163 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD4_DO1 | Digital out 1 (IO module 4) | 2 | RW | 2 Bytes Integer | 4163 | 34164 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD4_DO2 | Digital out 2 (IO module 4) | 2 | RW | 2 Bytes Integer | 4164 | 34165 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD4_DO3 | Digital out 3 (IO module 4) | 2 | RW | 2 Bytes Integer | 4165 | 34166 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALOUT_WRITE_LOGIC_CARD4_DO4 | Digital out 4 (IO module 4) | 2 | RW | 2 Bytes Integer | 4166 | 34167 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD1_1 | Analog out 1 (IO module 1) | 4 | RW | 4 Bytes Float | 4107 | 34108 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD1_2 | Analog out 2 (IO module 1) | 4 | RW | 4 Bytes Float | 4109 | 34110 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD1_3 | Analog out 3 (IO module 1) | 4 | RW | 4 Bytes Float | 4111 | 34112 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD1_4 | Analog out 4 (IO module 1) | 4 | RW | 4 Bytes Float | 4113 | 34114 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD2_1 | Analog out 1 (IO module 2) | 4 | RW | 4 Bytes Float | 4115 | 34116 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD2_2 | Analog out 2 (IO module 2) | 4 | RW | 4 Bytes Float | 4117 | 34118 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD2_3 | Analog out 3 (IO module 2) | 4 | RW | 4 Bytes Float | 4119 | 34120 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD2_4 | Analog out 4 (IO module 2) | 4 | RW | 4 Bytes Float | 4121 | 34122 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD3_1 | Analog out 1 (IO module 3) | 4 | RW | 4 Bytes Float | 4123 | 34124 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD3_2 | Analog out 2 (IO module 3) | 4 | RW | 4 Bytes Float | 4125 | 34126 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD3_3 | Analog out 3 (IO module 3) | 4 | RW | 4 Bytes Float | 4127 | 34128 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD3_4 | Analog out 4 (IO module 3) | 4 | RW | 4 Bytes Float | 4129 | 34130 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD4_1 | Analog out 1 (IO module 4) | 4 | RW | 4 Bytes Float | 4131 | 34132 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD4_2 | Analog out 1 (IO module 4) | 4 | RW | 4 Bytes Float | 4133 | 34134 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD4_3 | Analog out 1 (IO module 4) | 4 | RW | 4 Bytes Float | 4135 | 34136 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_OUT_CARD4_4 | Analog out 1 (IO module 4) | 4 | RW | 4 Bytes Float | 4137 | 34138 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD1_1 | Analog In 1 (IO module 1) , read only | 4 | R | 4 Bytes Float | 4043 | 34044 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD1_2 | Analog In 2 (IO module 1) , read only | 4 | R | 4 Bytes Float | 4045 | 34046 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD1_3 | Analog In 3 (IO module 1) , read only | 4 | R | 4 Bytes Float | 4047 | 34048 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD1_4 | Analog In 4 (IO module 1) , read only | 4 | R | 4 Bytes Float | 4049 | 34050 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD2_1 | Analog In 1 (IO module 2) , read only | 4 | R | 4 Bytes Float | 4051 | 34052 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD2_2 | Analog In 2 (IO module 2) , read only | 4 | R | 4 Bytes Float | 4053 | 34054 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD2_3 | Analog In 3 (IO module 2) , read only | 4 | R | 4 Bytes Float | 4055 | 34056 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD2_4 | Analog In 4 (IO module 2) , read only | 4 | R | 4 Bytes Float | 4057 | 34058 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD3_1 | Analog In 1 (IO module 3) , read only | 4 | R | 4 Bytes Float | 4059 | 34060 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD3_2 | Analog In 2 (IO module 3) , read only | 4 | R | 4 Bytes Float | 4061 | 34062 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD3_3 | Analog In 3 (IO module 3) , read only | 4 | R | 4 Bytes Float | 4063 | 34064 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD3_4 | Analog In 4 (IO module 3) , read only | 4 | R | 4 Bytes Float | 4065 | 34066 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD4_1 | Analog In 1 (IO module 4) , read only | 4 | R | 4 Bytes Float | 4067 | 34068 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD4_2 | Analog In 2 (IO module 4) , read only | 4 | R | 4 Bytes Float | 4069 | 34070 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD4_3 | Analog In 3 (IO module 4) , read only | 4 | R | 4 Bytes Float | 4071 | 34072 | Power Quality Analyzer G4400
PRM_CODE_MIO_ANALOG_IN_CARD4_4 | Analog In 4 (IO module 4) , read only | 4 | R | 4 Bytes Float | 4073 | 34074 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI1 | Logic bit D1 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4075 | 34076 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI2 | Logic bit D2 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4076 | 34077 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI3 | Logic bit D3 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4077 | 34078 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI4 | Logic bit D4 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4078 | 34079 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI5 | Logic bit D5 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4079 | 34080 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI6 | Logic bit D6 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4080 | 34081 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI7 | Logic bit D7 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4081 | 34082 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD1_DI8 | Logic bit D8 (io module 1) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4082 | 34083 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI1 | Logic bit D1 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4083 | 34084 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI2 | Logic bit D2 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4084 | 34085 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI3 | Logic bit D3 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4085 | 34086 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI4 | Logic bit D4 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4086 | 34087 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI5 | Logic bit D5 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4087 | 34088 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI6 | Logic bit D6 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4088 | 34089 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI7 | Logic bit D7 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4089 | 34090 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD2_DI8 | Logic bit D8 (io module 2) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4090 | 34091 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI1 | Logic bit D1 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4091 | 34092 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI2 | Logic bit D2 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4092 | 34093 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI3 | Logic bit D3 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4093 | 34094 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI4 | Logic bit D4 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4094 | 34095 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI5 | Logic bit D5 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4095 | 34096 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI6 | Logic bit D6 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4096 | 34097 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI7 | Logic bit D7 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4097 | 34098 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD3_DI8 | Logic bit D8 (io module 3) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4098 | 34099 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI1 | Logic bit D1 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4099 | 34100 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI2 | Logic bit D2 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4100 | 34101 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI3 | Logic bit D3 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4101 | 34102 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI4 | Logic bit D4 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4102 | 34103 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI5 | Logic bit D5 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4103 | 34104 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI6 | Logic bit D6 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4104 | 34105 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI7 | Logic bit D7 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4105 | 34106 | Power Quality Analyzer G4400
PRM_CODE_MIO_DIGITALIN_READ_LOGIC_CARD4_DI8 | Logic bit D8 (io module 4) sampled 16 times per cycle, read only | 2 | R | 2 Bytes Integer | 4106 | 34107 | Power Quality Analyzer G4400

## Harmonics & Interharmonics
Parameter | Description | MODBUS Code | Supported Device
PRM_CODE_V1_GROUP_HARMONICS | V1 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 592 | Power Quality Analyzer G4400
PRM_CODE_V2_GROUP_HARMONICS | V2 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 593 | Power Quality Analyzer G4400
PRM_CODE_V3_GROUP_HARMONICS | V3 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 594 | Power Quality Analyzer G4400
PRM_CODE_V12_GROUP_HARMONICS | V12 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 595 | Power Quality Analyzer G4400
PRM_CODE_V23_GROUP_HARMONICS | V23 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 596 | Power Quality Analyzer G4400
PRM_CODE_V31_GROUP_HARMONICS | V31 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 597 | Power Quality Analyzer G4400
PRM_CODE_I1_GROUP_HARMONICS | I1 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 598 | Power Quality Analyzer G4400
PRM_CODE_I2_GROUP_HARMONICS | I2 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 599 | Power Quality Analyzer G4400
PRM_CODE_I3_GROUP_HARMONICS | I3 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 600 | Power Quality Analyzer G4400
PRM_CODE_I12_GROUP_HARMONICS | I12 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 601 | Power Quality Analyzer G4400
PRM_CODE_I23_GROUP_HARMONICS | I23 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 602 | Power Quality Analyzer G4400
PRM_CODE_I31_GROUP_HARMONICS | I31 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 603 | Power Quality Analyzer G4400
PRM_CODE_VN_GROUP_HARMONICS | VN Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 604 | Power Quality Analyzer G4400
PRM_CODE_IN_GROUP_HARMONICS | IN Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 605 | Power Quality Analyzer G4400
PRM_CODE_IDC_GROUP_HARMONICS | IDC Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 970 | Power Quality Analyzer G4400
PRM_CODE_V1_GROUP_HARMONICS_MIN | V1 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1876 | Power Quality Analyzer G4400
PRM_CODE_V2_GROUP_HARMONICS_MIN | V2 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1877 | Power Quality Analyzer G4400
PRM_CODE_V3_GROUP_HARMONICS_MIN | V3 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1878 | Power Quality Analyzer G4400
PRM_CODE_V12_GROUP_HARMONICS_MIN | V12 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1879 | Power Quality Analyzer G4400
PRM_CODE_V23_GROUP_HARMONICS_MIN | V23 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1880 | Power Quality Analyzer G4400
PRM_CODE_V31_GROUP_HARMONICS_MIN | V31 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1881 | Power Quality Analyzer G4400
PRM_CODE_I1_GROUP_HARMONICS_MIN | I1 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1882 | Power Quality Analyzer G4400
PRM_CODE_I2_GROUP_HARMONICS_MIN | I2 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1883 | Power Quality Analyzer G4400
PRM_CODE_I3_GROUP_HARMONICS_MIN | I3 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1884 | Power Quality Analyzer G4400
PRM_CODE_I12_GROUP_HARMONICS_MIN | I12 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1885 | Power Quality Analyzer G4400
PRM_CODE_I23_GROUP_HARMONICS_MIN | I23 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1886 | Power Quality Analyzer G4400
PRM_CODE_I31_GROUP_HARMONICS_MIN | I31 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1887 | Power Quality Analyzer G4400
PRM_CODE_VN_GROUP_HARMONICS_MIN | VN Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1888 | Power Quality Analyzer G4400
PRM_CODE_IN_GROUP_HARMONICS_MIN | IN Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1889 | Power Quality Analyzer G4400
PRM_CODE_IDC_GROUP_HARMONICS_MIN | IDC Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1946 | Power Quality Analyzer G4400
PRM_CODE_V1_GROUP_HARMONICS_MAX | V1 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1890 | Power Quality Analyzer G4400
PRM_CODE_V2_GROUP_HARMONICS_MAX | V2 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1891 | Power Quality Analyzer G4400
PRM_CODE_V3_GROUP_HARMONICS_MAX | V3 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1892 | Power Quality Analyzer G4400
PRM_CODE_V12_GROUP_HARMONICS_MAX | V12 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1893 | Power Quality Analyzer G4400
PRM_CODE_V23_GROUP_HARMONICS_MAX | V23 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1894 | Power Quality Analyzer G4400
PRM_CODE_V31_GROUP_HARMONICS_MAX | V31 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1895 | Power Quality Analyzer G4400
PRM_CODE_I1_GROUP_HARMONICS_MAX | I1 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1896 | Power Quality Analyzer G4400
PRM_CODE_I2_GROUP_HARMONICS_MAX | I2 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1897 | Power Quality Analyzer G4400
PRM_CODE_I3_GROUP_HARMONICS_MAX | I3 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1898 | Power Quality Analyzer G4400
PRM_CODE_I12_GROUP_HARMONICS_MAX | I12 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1899 | Power Quality Analyzer G4400
PRM_CODE_I23_GROUP_HARMONICS_MAX | I23 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1900 | Power Quality Analyzer G4400
PRM_CODE_I31_GROUP_HARMONICS_MAX | I31 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1901 | Power Quality Analyzer G4400
PRM_CODE_VN_GROUP_HARMONICS_MAX | VN Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1902 | Power Quality Analyzer G4400
PRM_CODE_IN_GROUP_HARMONICS_MAX | IN Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1903 | Power Quality Analyzer G4400
PRM_CODE_IDC_GROUP_HARMONICS_MAX | IDC Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1947 | Power Quality Analyzer G4400
PRM_CODE_P1_GROUP_HARMONICS | P1 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1904 | Power Quality Analyzer G4400
PRM_CODE_P2_GROUP_HARMONICS | P2 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1905 | Power Quality Analyzer G4400
PRM_CODE_P3_GROUP_HARMONICS | P3 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1906 | Power Quality Analyzer G4400
PRM_CODE_Q1_GROUP_HARMONICS | Q1 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1910 | Power Quality Analyzer G4400
PRM_CODE_Q2_GROUP_HARMONICS | Q2 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1911 | Power Quality Analyzer G4400
PRM_CODE_Q3_GROUP_HARMONICS | Q3 Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1912 | Power Quality Analyzer G4400
PRM_CODE_PN_GROUP_HARMONICS | PN Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1916 | Power Quality Analyzer G4400
PRM_CODE_QN_GROUP_HARMONICS | QN Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1917 | Power Quality Analyzer G4400
PRM_CODE_P1_GROUP_HARMONICS_MIN | P1 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1918 | Power Quality Analyzer G4400
PRM_CODE_P2_GROUP_HARMONICS_MIN | P2 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1919 | Power Quality Analyzer G4400
PRM_CODE_P3_GROUP_HARMONICS_MIN | P3 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1920 | Power Quality Analyzer G4400
PRM_CODE_Q1_GROUP_HARMONICS_MIN | Q1 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1924 | Power Quality Analyzer G4400
PRM_CODE_Q2_GROUP_HARMONICS_MIN | Q2 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1925 | Power Quality Analyzer G4400
PRM_CODE_Q3_GROUP_HARMONICS_MIN | Q3 Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1926 | Power Quality Analyzer G4400
PRM_CODE_PN_GROUP_HARMONICS_MIN | PN Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1930 | Power Quality Analyzer G4400
PRM_CODE_QN_GROUP_HARMONICS_MIN | QN Minimum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1931 | Power Quality Analyzer G4400
PRM_CODE_P1_GROUP_HARMONICS_MAX | P1 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1932 | Power Quality Analyzer G4400
PRM_CODE_P2_GROUP_HARMONICS_MAX | P2 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1933 | Power Quality Analyzer G4400
PRM_CODE_P3_GROUP_HARMONICS_MAX | P3 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1934 | Power Quality Analyzer G4400
PRM_CODE_Q1_GROUP_HARMONICS_MAX | Q1 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1938 | Power Quality Analyzer G4400
PRM_CODE_Q2_GROUP_HARMONICS_MAX | Q2 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1939 | Power Quality Analyzer G4400
PRM_CODE_Q3_GROUP_HARMONICS_MAX | Q3 Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1940 | Power Quality Analyzer G4400
PRM_CODE_PN_GROUP_HARMONICS_MAX | PN Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1944 | Power Quality Analyzer G4400
PRM_CODE_QN_GROUP_HARMONICS_MAX | QN Maximum Harmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 1945 | Power Quality Analyzer G4400
PRM_CODE_V1_PHASES_GROUP_HARMONICS | V1 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2329 | Power Quality Analyzer G4400
PRM_CODE_V2_PHASES_GROUP_HARMONICS | V2 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2330 | Power Quality Analyzer G4400
PRM_CODE_V3_PHASES_GROUP_HARMONICS | V3 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2331 | Power Quality Analyzer G4400
PRM_CODE_V12_PHASES_GROUP_HARMONICS | V12 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2332 | Power Quality Analyzer G4400
PRM_CODE_V23_PHASES_GROUP_HARMONICS | V23 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2333 | Power Quality Analyzer G4400
PRM_CODE_V31_PHASES_GROUP_HARMONICS | V31 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2334 | Power Quality Analyzer G4400
PRM_CODE_I1_PHASES_GROUP_HARMONICS | I1 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2335 | Power Quality Analyzer G4400
PRM_CODE_I2_PHASES_GROUP_HARMONICS | I2 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2336 | Power Quality Analyzer G4400
PRM_CODE_I3_PHASES_GROUP_HARMONICS | I3 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2337 | Power Quality Analyzer G4400
PRM_CODE_I12_PHASES_GROUP_HARMONICS | I12 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2338 | Power Quality Analyzer G4400
PRM_CODE_I23_PHASES_GROUP_HARMONICS | I23 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2339 | Power Quality Analyzer G4400
PRM_CODE_I31_PHASES_GROUP_HARMONICS | I31 Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2340 | Power Quality Analyzer G4400
PRM_CODE_VN_PHASES_GROUP_HARMONICS | VN Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2341 | Power Quality Analyzer G4400
PRM_CODE_IN_PHASES_GROUP_HARMONICS | IN Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2342 | Power Quality Analyzer G4400
PRM_CODE_IDC_PHASES_GROUP_HARMONICS | IDC Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2343 | Power Quality Analyzer G4400
PRM_CODE_V1_PHASES_GROUP_HARMONICS_MIN | V1 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2344 | Power Quality Analyzer G4400
PRM_CODE_V2_PHASES_GROUP_HARMONICS_MIN | V2 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2345 | Power Quality Analyzer G4400
PRM_CODE_V3_PHASES_GROUP_HARMONICS_MIN | V3 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2346 | Power Quality Analyzer G4400
PRM_CODE_V12_PHASES_GROUP_HARMONICS_MIN | V12 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2347 | Power Quality Analyzer G4400
PRM_CODE_V23_PHASES_GROUP_HARMONICS_MIN | V23 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2348 | Power Quality Analyzer G4400
PRM_CODE_V31_PHASES_GROUP_HARMONICS_MIN | V31 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2349 | Power Quality Analyzer G4400
PRM_CODE_I1_PHASES_GROUP_HARMONICS_MIN | I1 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2350 | Power Quality Analyzer G4400
PRM_CODE_I2_PHASES_GROUP_HARMONICS_MIN | I2 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2351 | Power Quality Analyzer G4400
PRM_CODE_I3_PHASES_GROUP_HARMONICS_MIN | I3 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2352 | Power Quality Analyzer G4400
PRM_CODE_I12_PHASES_GROUP_HARMONICS_MIN | I12 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2353 | Power Quality Analyzer G4400
PRM_CODE_I23_PHASES_GROUP_HARMONICS_MIN | I23 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2354 | Power Quality Analyzer G4400
PRM_CODE_I31_PHASES_GROUP_HARMONICS_MIN | I31 Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2355 | Power Quality Analyzer G4400
PRM_CODE_VN_PHASES_GROUP_HARMONICS_MIN | VN Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2356 | Power Quality Analyzer G4400
PRM_CODE_IN_PHASES_GROUP_HARMONICS_MIN | IN Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2357 | Power Quality Analyzer G4400
PRM_CODE_IDC_PHASES_GROUP_HARMONICS_MIN | IDC Minimum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2358 | Power Quality Analyzer G4400
PRM_CODE_V1_PHASES_GROUP_HARMONICS_MAX | V1 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2359 | Power Quality Analyzer G4400
PRM_CODE_V2_PHASES_GROUP_HARMONICS_MAX | V2 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2360 | Power Quality Analyzer G4400
PRM_CODE_V3_PHASES_GROUP_HARMONICS_MAX | V3 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2361 | Power Quality Analyzer G4400
PRM_CODE_V12_PHASES_GROUP_HARMONICS_MAX | V12 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2362 | Power Quality Analyzer G4400
PRM_CODE_V23_PHASES_GROUP_HARMONICS_MAX | V23 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2363 | Power Quality Analyzer G4400
PRM_CODE_V31_PHASES_GROUP_HARMONICS_MAX | V31 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2364 | Power Quality Analyzer G4400
PRM_CODE_I1_PHASES_GROUP_HARMONICS_MAX | I1 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2365 | Power Quality Analyzer G4400
PRM_CODE_I2_PHASES_GROUP_HARMONICS_MAX | I2 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2366 | Power Quality Analyzer G4400
PRM_CODE_I3_PHASES_GROUP_HARMONICS_MAX | I3 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2367 | Power Quality Analyzer G4400
PRM_CODE_I12_PHASES_GROUP_HARMONICS_MAX | I12 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2368 | Power Quality Analyzer G4400
PRM_CODE_I23_PHASES_GROUP_HARMONICS_MAX | I23 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2369 | Power Quality Analyzer G4400
PRM_CODE_I31_PHASES_GROUP_HARMONICS_MAX | I31 Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2370 | Power Quality Analyzer G4400
PRM_CODE_VN_PHASES_GROUP_HARMONICS_MAX | VN Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2371 | Power Quality Analyzer G4400
PRM_CODE_IN_PHASES_GROUP_HARMONICS_MAX | IN Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2372 | Power Quality Analyzer G4400
PRM_CODE_IDC_PHASES_GROUP_HARMONICS_MAX | IDC Maximum Harmonic (Phasor) # 1-50 th, calculated over 10/12 cycles | 2373 | Power Quality Analyzer G4400
PRM_CODE_V1_INTERHARMONIC_GROUP_HARMONICS | V1 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2374 | Power Quality Analyzer G4400
PRM_CODE_V2_INTERHARMONIC_GROUP_HARMONICS | V2 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2375 | Power Quality Analyzer G4400
PRM_CODE_V3_INTERHARMONIC_GROUP_HARMONICS | V3 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2376 | Power Quality Analyzer G4400
PRM_CODE_V12_INTERHARMONIC_GROUP_HARMONICS | V12 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2377 | Power Quality Analyzer G4400
PRM_CODE_V23_INTERHARMONIC_GROUP_HARMONICS | V23 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2378 | Power Quality Analyzer G4400
PRM_CODE_V31_INTERHARMONIC_GROUP_HARMONICS | V31 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2379 | Power Quality Analyzer G4400
PRM_CODE_I1_INTERHARMONIC_GROUP_HARMONICS | I1 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2380 | Power Quality Analyzer G4400
PRM_CODE_I2_INTERHARMONIC_GROUP_HARMONICS | I2 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2381 | Power Quality Analyzer G4400
PRM_CODE_I3_INTERHARMONIC_GROUP_HARMONICS | I3 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2382 | Power Quality Analyzer G4400
PRM_CODE_I12_INTERHARMONIC_GROUP_HARMONICS | I12 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2383 | Power Quality Analyzer G4400
PRM_CODE_I23_INTERHARMONIC_GROUP_HARMONICS | I23 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2384 | Power Quality Analyzer G4400
PRM_CODE_I31_INTERHARMONIC_GROUP_HARMONICS | I31 Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2385 | Power Quality Analyzer G4400
PRM_CODE_VN_INTERHARMONIC_GROUP_HARMONICS | VN Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2386 | Power Quality Analyzer G4400
PRM_CODE_IN_INTERHARMONIC_GROUP_HARMONICS | IN Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2387 | Power Quality Analyzer G4400
PRM_CODE_IDC_INTERHARMONIC_GROUP_HARMONICS | IDC Interharmonic (Amplitude) # 1-50 th, calculated over 10/12 cycles | 2388 | Power Quality Analyzer G4400
PRM_CODE_V1_SUBGROUP_HARMONICS | V1 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2389 | Power Quality Analyzer G4400
PRM_CODE_V2_SUBGROUP_HARMONICS | V2 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2390 | Power Quality Analyzer G4400
PRM_CODE_V3_SUBGROUP_HARMONICS | V3 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2391 | Power Quality Analyzer G4400
PRM_CODE_V12_SUBGROUP_HARMONICS | V12 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2392 | Power Quality Analyzer G4400
PRM_CODE_V23_SUBGROUP_HARMONICS | V23 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2393 | Power Quality Analyzer G4400
PRM_CODE_V31_SUBGROUP_HARMONICS | V31 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2394 | Power Quality Analyzer G4400
PRM_CODE_I1_SUBGROUP_HARMONICS | I1 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2395 | Power Quality Analyzer G4400
PRM_CODE_I2_SUBGROUP_HARMONICS | I2 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2396 | Power Quality Analyzer G4400
PRM_CODE_I3_SUBGROUP_HARMONICS | I3 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2397 | Power Quality Analyzer G4400
PRM_CODE_I12_SUBGROUP_HARMONICS | I12 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2398 | Power Quality Analyzer G4400
PRM_CODE_I23_SUBGROUP_HARMONICS | I23 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2399 | Power Quality Analyzer G4400
PRM_CODE_I31_SUBGROUP_HARMONICS | I31 Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2400 | Power Quality Analyzer G4400
PRM_CODE_VN_SUBGROUP_HARMONICS | VN Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2401 | Power Quality Analyzer G4400
PRM_CODE_IN_SUBGROUP_HARMONICS | IN Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2402 | Power Quality Analyzer G4400
PRM_CODE_IDC_SUBGROUP_HARMONICS | IDC Subgroup Harmonics # 1-50 th, calculated over 10/12 cycles | 2403 | Power Quality Analyzer G4400
PRM_CODE_V1_INTERHARMONIC_SUBGROUP | V1 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2404 | Power Quality Analyzer G4400
PRM_CODE_V2_INTERHARMONIC_SUBGROUP | V2 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2405 | Power Quality Analyzer G4400
PRM_CODE_V3_INTERHARMONIC_SUBGROUP | V3 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2406 | Power Quality Analyzer G4400
PRM_CODE_V12_INTERHARMONIC_SUBGROUP | V12 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2407 | Power Quality Analyzer G4400
PRM_CODE_V23_INTERHARMONIC_SUBGROUP | V23 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2408 | Power Quality Analyzer G4400
PRM_CODE_V31_INTERHARMONIC_SUBGROUP | V31 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2409 | Power Quality Analyzer G4400
PRM_CODE_I1_INTERHARMONIC_SUBGROUP | I1 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2410 | Power Quality Analyzer G4400
PRM_CODE_I2_INTERHARMONIC_SUBGROUP | I2 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2411 | Power Quality Analyzer G4400
PRM_CODE_I3_INTERHARMONIC_SUBGROUP | I3 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2412 | Power Quality Analyzer G4400
PRM_CODE_I12_INTERHARMONIC_SUBGROUP | I12 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2413 | Power Quality Analyzer G4400
PRM_CODE_I23_INTERHARMONIC_SUBGROUP | I23 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2414 | Power Quality Analyzer G4400
PRM_CODE_I31_INTERHARMONIC_SUBGROUP | I31 Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2415 | Power Quality Analyzer G4400
PRM_CODE_VN_INTERHARMONIC_SUBGROUP | VN Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2416 | Power Quality Analyzer G4400
PRM_CODE_IN_INTERHARMONIC_SUBGROUP | IN Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2417 | Power Quality Analyzer G4400
PRM_CODE_IDC_INTERHARMONIC_SUBGROUP | IDC Interharmonic subgroup # 1-50 th, calculated over 10/12 cycles | 2418 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_VN_SUBGROUP_HARMONICS | VN Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6340 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V1_SUBGROUP_HARMONICS | V1 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6341 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V2_SUBGROUP_HARMONICS | V2 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6342 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V3_SUBGROUP_HARMONICS | V3 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6343 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V12_SUBGROUP_HARMONICS | V12 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6344 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V23_SUBGROUP_HARMONICS | V23 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6345 | Power Quality Analyzer G4400
PRM_CODE_AVG_LONG_2HOUR_V31_SUBGROUP_HARMONICS | V31 Average Subgroup harmonics # 1-50 th, calculated over 2 hours | 6346 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_VN_SUBGROUP_HARMONICS | VN Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6277 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V1_SUBGROUP_HARMONICS | V1 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6278 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V2_SUBGROUP_HARMONICS | V2 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6279 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V3_SUBGROUP_HARMONICS | V3 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6280 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V12_SUBGROUP_HARMONICS | V12 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6281 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V23_SUBGROUP_HARMONICS | V23 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6282 | Power Quality Analyzer G4400
PRM_CODE_AVG_3SEC_V31_SUBGROUP_HARMONICS | V31 Average Subgroup Harmonics # 1-50 th, calculated over 3 seconds | 6283 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_VN_SUBGROUP_HARMONICS | VN Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6333 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V1_SUBGROUP_HARMONICS | V1 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6334 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V2_SUBGROUP_HARMONICS | V2 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6335 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V3_SUBGROUP_HARMONICS | V3 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6336 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V12_SUBGROUP_HARMONICS | V12 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6337 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V23_SUBGROUP_HARMONICS | V23 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6338 | Power Quality Analyzer G4400
PRM_CODE_AVG_10MIN_V31_SUBGROUP_HARMONICS | V31 Average Subgroup Harmonics # 1-50 th, calculated over 10 minutes | 6339 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_3SEC_SUBGROUP_INTERHARMONICS | V1 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1993 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_3SEC_SUBGROUP_INTERHARMONICS | V2 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1994 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_3SEC_SUBGROUP_INTERHARMONICS | V3 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1995 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_3SEC_SUBGROUP_INTERHARMONICS | V12 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1996 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_3SEC_SUBGROUP_INTERHARMONICS | V23 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1997 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_3SEC_SUBGROUP_INTERHARMONICS | V31 Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1998 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_3SEC_SUBGROUP_INTERHARMONICS | VN Average Subgroup Interharmonics # 1-50 th, calculated over 3 seconds | 1999 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_10MIN_SUBGROUP_INTERHARMONICS | V1 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2000 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_10MIN_SUBGROUP_INTERHARMONICS | V2 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2001 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_10MIN_SUBGROUP_INTERHARMONICS | V3 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2002 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_10MIN_SUBGROUP_INTERHARMONICS | V12 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2003 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_10MIN_SUBGROUP_INTERHARMONICS | V23 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2004 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_10MIN_SUBGROUP_INTERHARMONICS | V31 Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2005 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_10MIN_SUBGROUP_INTERHARMONICS | VN Average Subgroup Interharmonics # 1-50 th, calculated over 10 minutes | 2006 | Power Quality Analyzer G4400
PRM_CODE_LONG_V1_2HOURS_SUBGROUP_INTERHARMONICS | V1 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2007 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_2HOURS_SUBGROUP_INTERHARMONICS | V2 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2008 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_2HOURS_SUBGROUP_INTERHARMONICS | V3 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2009 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_2HOURS_SUBGROUP_INTERHARMONICS | V12 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2010 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_2HOURS_SUBGROUP_INTERHARMONICS | V23 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2011 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_2HOURS_SUBGROUP_INTERHARMONICS | V31 Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2012 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_2HOURS_SUBGROUP_INTERHARMONICS | VN Average Subgroup Interharmonics # 1-50 th, calculated over 2 hours | 2013 | Power Quality Analyzer G4400

## Spectrum
Parameter | Description | MODBUS Code | Supported Device
PRM_CODE_LONG_V1_FFT_AMP | V1's Spectrum Amplitude | 1978 | Power Quality Analyzer G4400
PRM_CODE_LONG_V2_FFT_AMP | V2's Spectrum Amplitude | 1979 | Power Quality Analyzer G4400
PRM_CODE_LONG_V3_FFT_AMP | V3's Spectrum Amplitude | 1980 | Power Quality Analyzer G4400
PRM_CODE_LONG_V12_FFT_AMP | V12's Spectrum Amplitude | 1981 | Power Quality Analyzer G4400
PRM_CODE_LONG_V23_FFT_AMP | V23's Spectrum Amplitude | 1982 | Power Quality Analyzer G4400
PRM_CODE_LONG_V31_FFT_AMP | V31's Spectrum Amplitude | 1983 | Power Quality Analyzer G4400
PRM_CODE_LONG_I1_FFT_AMP | I1's Spectrum Amplitude | 1984 | Power Quality Analyzer G4400
PRM_CODE_LONG_I2_FFT_AMP | I2's Spectrum Amplitude | 1985 | Power Quality Analyzer G4400
PRM_CODE_LONG_I3_FFT_AMP | I3's Spectrum Amplitude | 1986 | Power Quality Analyzer G4400
PRM_CODE_LONG_I12_FFT_AMP | I12's Spectrum Amplitude | 1987 | Power Quality Analyzer G4400
PRM_CODE_LONG_I23_FFT_AMP | I23's Spectrum Amplitude | 1988 | Power Quality Analyzer G4400
PRM_CODE_LONG_I31_FFT_AMP | I31's Spectrum Amplitude | 1989 | Power Quality Analyzer G4400
PRM_CODE_LONG_VN_FFT_AMP | VN's Spectrum Amplitude | 1990 | Power Quality Analyzer G4400
PRM_CODE_LONG_IN_FFT_AMP | IN's Spectrum Amplitude | 1991 | Power Quality Analyzer G4400
PRM_CODE_LONG_IDC_FFT_AMP | IDC's Spectrum Amplitude | 1992 | Power Quality Analyzer G4400